## Items for multiple tests
Master Form 101359 Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TODO ID
| Thermometers-Working |  |  | annually |  | see separate list |  |  |  | 1
| Thermometer- Beta Probe |  |  | every 2 years |  | 2014-03-19 | Outside Vendor | 2016-03-19 |  | 2
| Thermometer- Beta Probe |  |  | every 2 years | 3308115 | 2016-01-21 | Outside Vendor | 2018-01-21 |  | 3
| Hydrometers |  |  | annually | [5 empty] | 4
| Tachometer |  |  | annually | Centech | 2016-08-31 | Outside Vendor Micro Precision Calibration) | 2017-08-31 |  | 5
| Tachometer |  |  | annually | Pyle Laser | 2016-05-25 | MS | 2017-05-25 |  | 6
| Barometer |  |  | per NIST Traceable certificate | [5 empty] | 7
| calibration weights | E617 | Recalibration | annually |  | 2016-02-28 | Outside Vendor | 2017-02-28 |  | 8
| Viscometer | D445 | Verification | Annually |  | see separate list |  |  |  | 9
| Timers |  | Verification | Annually |  | see separate list |  |  |  | 10
| flow meter |  | Calibration | annually | RE100548 | 2016-08-31 | Outside Vendor Micro Precision Calibration) | 2017-08-31 |  | 11
| micropipetters |  |  | Monthly |  | see separate list |  |  |  | 12
| Thermocouple Calibrator |  | Calibration | annually | 11000172 | 2016-05-26 | MS | 2017-05-26 |  | 13
| Digital Calipers |  |  | Annually | Mitutoyo 11408222 | 2016-02-03 | Aldinger | 2017-02-03 |  | 14
| Radiation Alert |  | Calibration | annually | 35224 | 2016-03-25 | S.E. International, Inc. | 2017-03-25 |  | 15
| 500 PSI Test Gauge |  | Calibration | annually | 211H16560009 | 2016-09-15 | Additel | 2017-09-15 |  | 16
| 10,000 psi test gauge |  | Calibration | annually | 211H166F0001 | 2016-09-15 | Additel | 2017-09-15 |  | 17
| 30 psi test gauge |  | Calibration | annually | Ashcroft | 2016-11-14 | MS | 2017-11-14 |  | 18

## Specifications
Calibration, Verification, Standardization Schedule
ASTM | Requirement | Current ASTM Version | TODO ID
D975 | Diesel | D975-15c | 1
D4814 | Gas | D4814-16d | 2
D4806 | Ethanol | D4806-16a | 3
D1655 | Jet A | D1655-16a | 4
D6751 | Biodiesel | D6751-15ce1 | 5
D6299 | SQC | D6299-16 | 6
D446 | Viscometers | D446-12 | 7
E77 | Thermometers | E77-14 | 8
E126 | Hydrometers | E126-13a | 9
E617 | Weights | E617-13 | 10
E29 | Rounding | E29-13 | 11
D6300 | Precision and Bias | D6300-16a | 12

## BiodieselGC Room
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| FTIR | D7371 | PLS Calibration -see Annex A1 | Initial and per SQC | A2091 | 2014-09-22 | BB | per SQC |  |  |  | 4000.20.092 10/28/14 | D7371-14 | 1
|  |  | Qualification- SEQ<PSEQ | After calibration | [9 empty] | 2
|  | E2412 | [10 empty] | E2412-10 | 3
| Perkin Elmer FTIR | D7371 | PLS Calibration -see Annex A1 | Initial and per SQC | C90420 | 2016-08-18 | JM | per SQC |  |  |  | 4000.20.092 10/28/14 | D7371-14
|  |  | Qualification- SEQ<PSEQ | After calibration
|  | E2412 | [10 empty] | E2412-10
| Lawler Oxidation Stability | EN15751 | Air Flow 10 + 1 L/h | annually | Lawler 321-1 | 2016-03-11 | SR | 2017-03-11 |  | 2013-11-07 | 2016-11-07 | 4000.20.081 11/8/13 | EN15751
|  |  | temp of block 110C | annually | Lawler 321-1 | 2016-03-11 | SR | 2017-03-11
| GC-FID | D5501 | Calibration | on setup and as required by SQC | US00005993 | 2015-03-07 | SR | per SQC |  | 2014-06-02 | 2017-06-02 | 4000.20.014 6/14/11 | D5501-12e1
|  | D5501 | Calibration | on setup and as required by SQC | LR47359C | 2015-03-07 | SR | per SQC
| GC-FID | D6584 | Calibration | on setup and as required by SQC | N6299 | 2015-05-12 | SR | as required by SQC |  | 2013-07-08 | 2016-07-08 | 4000.20.021 7/24/13 | D6584-13e1
Front Lab | Petrospec | D6277 | per Annex 1 of method | per SQC | 5600308 | 2015-05-11 | PAC | per SQC |  | 2012-12-05 | 2015-12-05 | 4000.20.054 5/25/13 | D6277-07(2012)
|  | D5845 | per Annex 1 of method | per SQC | 5600308 | 2015-05-11 | PAC | per SQC |  | 2012-12-28 | 2015-12-28 | 4000.20.058 3/1/13 | D5845-02(2011)
| Grabner | D5191 | Pressure Transducer | every six months | 07-01-1136 |  | Done by PSSC |  |  | 2014-07-01 | 2017-07-01 | 4000.20.023 7/2/14 | D5191-13
|  |  | Thermometer | every six months | 07-01-1136 |  | Done by PSSC
| RVP Grabner | D5191 | Pressure Transducer | every six months | 20-210-2829 | 2016-08-04 | Done by PSSC | 2017-02-04 |  | 2014-07-01 | Due- Method updated | 4000.20.023 7/2/14 | D5191-15
|  |  | Thermometer | every six months | 20-210-2829 | 2016-08-04 | Done by PSSC | 2017-02-04
|  | D5191 | Pressure Transducer | every six months | 07-01-1136 | 2016-09-23 | Done by PSSC | 2017-03-23
|  |  | Thermometer | every six months | 07-01-1136 | 2016-09-23 | Done by PSSC | 2017-03-23
|  | D3227 | Standardize AgNO3 | Weekly | Titrando 809 SN10809001013164 | [7 empty] | D3227-13
| Omax M83E5 | FTM3005 | Calibrate camera at 10X and 40X | Annually | G016233378 | 2016-09-15 | BTB | 2017-09-15
| Perkin Elmer Claris 500 GC | D2887 | Calibrate boiling point curve | As needed per SQC | 650509093010 | 2016-10-28 | HL | per SQC

## Front Lab
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| pHe | D6423 | Standardize | daily | 925418 | per use |  |  |  | 2012-12-06 | Due- Method updated | 4000.20.034 6/15/11 | D6423-14 | 1
| Hylite | D7463 | RLU background/control <=0.01 | annually | 0012975J06 | 2016-03-17 | SR | 2017-03-17 |  | 2012-12-27 | 2015-12-27 | 4000.20.041 3/6/13 | D7463-15 | 2
| LuminUltra | D7687 | RLU background/control <=0.01 | annually | 0012975J06 | 2016-03-17 | SR | 2017-03-17 |  | None |  | 4000.20.077 9/30/13 | D7687-11 | 3
| Karl Fischer Columetric | E1064 | refer to instrument manual |  | 081610 | 2016-10-17 | SC | 2016-12-17 |  | 2012-12-06 | 2015-12-06 |  | E1064-12 | 4
| Karl Fischer Columetric | D6304 | Water Standard to check response | every 10 samples and with new reagent | 081610 | 2016-11-17 | JM | 2017-01-17 |  | 2013-06-03 | 2016-06-03 | 4000.20.039 7/24/11 and 4000.20.004 10/21/13 | D6304-07 (withdrawn 2016) | 5
| Karl Fischer Columetric | D1744 | Standardize reagent | daily | 081610 | [7 empty] | D1744-13 | 6
| Karl Fischer Columetric | D4377 | Standardize reagent | daily | 081610 | [7 empty] | D4377-00 (2011) | 7
| Karl Fischer Volumetric | D4377 | Standardize reagent | daily | 8.915.80002EPN | [7 empty] | D4377-00 (2011) | 8
| Karl Fischer Volumetric | E203 | standardize with Hydranal | every 10 samples and with new reagent | 8.915.80002EPN | 2016-11-15 | JM | 2016-11-22 |  | 2014-06-02 | 2017-06-02 | 4000.20.063 5/20/13 | E203-08 | 9
|  |  | SQC check to verify | daily | 8.915.80002EPN | with use | [7 empty] | 10
| Smoke Point | D1322 | Calibrate with 2 standards per analyst | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 4/4/14 (741.8 mm Hg) | SR | pressure less than 735.8 or higher than 746.3 |  | 2014-06-02 | 2015-06-02 | 4000.20.018 6/14/11 | D1322-12e2 | 11
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 4/17/14 (738.8 mm Hg) | JA | pressure less than 741.81 or higher than 752.31 | [5 empty] | 12
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 7/24/14 (744.8 mm Hg) | SA | pressure less than 739.55 or higher than 750.05 | [5 empty] | 13
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 7/24/14 (744.8 mm Hg) | RD | pressure less than 739.55 or higher than 750.05 | [5 empty] | 14
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 7/24/14 (744.8 mm Hg) | JL | pressure less than 739.55 or higher than 750.05 | [5 empty] | 15
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 7/24/14 (744.8 mm Hg) | SI | pressure less than 739.55 or higher than 750.05 | [5 empty] | 16
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 2/5/15 (754.5 mm Hg) | MC | pressure less than 749.25 or higher than 759.75 | [5 empty] | 17
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Koehler | 2/5/15 (748.6 mm Hg) | AG | pressure less than 743.35 or higher than 753.85 | [5 empty] | 18
| Smoke Point | D1322 | Calibrate with 2 standards per analyst | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29431 | 4/4/14 (741.05 mm Hg) and 4/7/15 (747.06 mm Hg) | SR | pressure less than 735.8 or higher than 752.31 | SR | 2015-09-16 | 2016-09-16 | 4000.20.018 6/14/11 | D1322-15e1 | 19
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | 539401 | 4/7/2014 (747.06 mm Hg) | JA | pressure less than 741.81 or higher than 752.31 | [5 empty] | 20
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29431 | 7/24/14 (744.8 mm Hg) | SA | pressure less than 739.55 or higher than 750.05 | [5 empty] | 21
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29432 | 7/24/14 (744.8 mm Hg) | RD | pressure less than 739.55 or higher than 750.05 | [5 empty] | 22
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | 539402 | 7/24/14 (744.8 mm Hg) | JL | pressure less than 739.55 or higher than 750.05 | [5 empty] | 23
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29432 | 2016-08-05 | AB | pressure less than 738.81 or higher than 749.3 | [5 empty] | 24
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29433 | 11/13/2015 (741.06 mm HG) | NL | pressure less than 735.81 or higher than 746.31 | [5 empty] | 25
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | 539403 | 11/13/2015 (741.06 mm HG) | NL | pressure less than 735.81 or higher than 746.32 | [5 empty] | 26
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29433 | 11/13/2015 (741.06 mm HG) | NL | pressure less than 735.81 or higher than 746.33 | [5 empty] | 27
| Gum Bath | D381 | Air flow 600 + 90 mL/s at ambient temp | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29434 | 11/13/2015 (741.06 mm HG) | NL | pressure less than 735.81 or higher than 746.34 | Moved to Floresville | 2014-06-03 | 2015-06-03 | 4000.20.022 3/30/14 | D381-12 | 28
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | 539404 | 6/23/15 (747.81 mm HG) | JB | pressure less than 742.56 or higher than 753.06 | [5 empty] | 29
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29434 | 2/5/2015 (754.5 mm HG) | MC | pressure less than 749.25 or higher than 759.75 | [5 empty] | 30
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29435 | 4/20/2016 (747.06 mm HG) | JS | pressure less than 741.81 or higher than 752.31 | [5 empty] | 31
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29435 | 2016-08-05 | JM | pressure less than 738.81 or higher than 749.3 | [5 empty] | 32
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29435 | 6/17/2015 (741.8 mm HG) | SC | pressure less than 736.55 or higher than 747.05 | [5 empty] | 33
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29435 | 2016-08-12 | RL | pressure less than 735.06 or higher than 745.56 | [5 empty] | 34
|  |  |  | Change of 0.7 kPa (5.25 mm Hg)for barometric pressure | Permian 29435 | 5/24/16 (744.06 mm HG) | MM | pressure less than 738.80 or higher than 749.31 | [5 empty] | 35
|  |  | Temperature of bath for gas/ ethanol 160-165C | annually | HGT 915 069151085 | 2014-03-25 | EA | 2015-03-25 | Moved to Floresville |  |  |  |  | 36
|  |  | Temperature of bath for Jet 232-246C | annually | HGT 915 069151085 | 2014-03-25 | EA | 2015-03-25 | Moved to Floresville |  |  |  |  | 37
| Gum Bath | D381 | Air flow 600 + 90 mL/s at ambient temp | annually | Precision Bath #1 | 2016-04-11 | SR | 2017-04-11 |  | 2015-08-26 | 2016-08-26 | 4000.20.022 3/30/14 | D381-12 | 38
|  |  | Temperature of bath for gas/ ethanol 160-165C | annually | Precision Bath #1 | 2016-04-11 | SR | 2017-04-11 | [5 empty] | 39
|  |  | Temperature of bath for Jet 232-246C | annually | Precision Bath #1 | 2016-04-11 | SR | 2017-04-11 | [5 empty] | 40
|  |  | Air flow 600 + 90 mL/s at ambient temp | annually | Precision Bath #2 | 2016-08-29 | SR | 2017-08-29 | [5 empty] | 41
|  |  | Temperature of bath for gas/ ethanol 160-165C | annually | Precision Bath #2 | 2016-08-29 | SR | 2017-08-29 | [5 empty] | 42
|  |  | Temperature of bath for Jet 232-246C | annually | Precision Bath #2 | 2016-08-29 | SR | 2017-08-29 | [5 empty] | 43
| ICP | D1688M | Calibrate with multiple standards | daily | 8003785 | with use |  |  |  | 2013-06-30 | 2016-06-30 | 4000.20.036 7/2/13 | d1688-12 | 44
|  | D7111 | 2 point calibration | daily | 8003785 | with use |  |  |  | 2013-07-04 | Method updated | 4000.20.066 1/29/14 | D7111-15a | 45
|  |  | Method Detection Limit Study | daily as required by client reporting | 8003785 | with use | [7 empty] | 46
|  | D5185 | Linear Range Study | on setup | 8003785 | 2013-07-13 |  |  |  | 2013-07-04 | 2016-07-04 | 4000.20.066 1/29/14 | D5185-13e1 | 47
|  |  | 2 point calibration | daily | 8003785 | with use | [7 empty] | 48
|  | D7691 | Linear Range Study | on setup | 8003785 | 2013-07-13 |  |  |  | 2013-07-05 | 2016-07-04 | 4000.20.066 1/29/14 | D7691-11e1 | 49
|  |  | 2 point calibration | daily | 8003785 | with use | [7 empty] | 50
|  | D4951 | 2 point calibration | daily | 8003785 | with use |  |  |  | 2013-07-04 | Method updated | 4000.20.066 1/29/14 | D4951-14 | 51
|  | EN 14538 | 5 point calibration. R>=0.995 | daily | 8003785 | with use |  |  |  | 2013-07-02 | 2016-07-04 | 4000.20.066 1/29/14 | EN14538 | 52
|  | D5708 | 2 point calibration | daily | 8003785 | with use |  |  |  | 2013-07-05 | Method updated | 4000.20.066 1/29/14 | D5708-15 | 53
|  | IP 501 | 5 point calibration. R>=0.995 | daily | 8003785 | with use | [6 empty] | IP 501 | 54
|  | EPA 6010 | 2 point calibration | daily | 8003785 | with use | [6 empty] | EPA 6010D | 55
|  | UOP 389 | 5 point calibration. R>=0.995 | daily | 8003785 | with use |  |  |  | 2015-06-07 | 2018-06-07 | 4000.20.099 6/4/15 | UOP 389-14 | 56
|  | MIL | 5 point calibration. R>=0.995, 5-replicates per standard and sample | daily | 8003785 | with use | [7 empty] | 57
| ICP | D1688M | Calibrate with multiple standards | daily | PE Optima 8000 | with use |  |  |  | 2013-06-30 | 2016-06-30 | 4000.20.036 7/2/13 | d1688-12 | 58
|  | D7111 | 2 point calibration | daily | PE Optima 8000 | with use |  |  |  | 2013-07-04 | Method updated | 4000.20.066 1/29/14 | D7111-15a | 59
|  |  | Method Detection Limit Study | daily as required by client reporting | PE Optima 8000 | with use | [7 empty] | 60
|  | D5185 | Linear Range Study | on setup | PE Optima 8000 |  |  |  |  | 2013-07-04 | 2016-07-04 | 4000.20.066 1/29/14 | D5185-13e1 | 61
|  |  | 2 point calibration | daily | PE Optima 8000 | with use | [7 empty] | 62
|  | D7691 | Linear Range Study | on setup | PE Optima 8000 |  |  |  |  | 2013-07-05 | 2016-07-04 | 4000.20.066 1/29/14 | D7691-11e1 | 63
|  |  | 2 point calibration | daily | PE Optima 8000 | with use | [7 empty] | 64
|  | D4951 | 2 point calibration | daily | PE Optima 8000 | with use |  |  |  | 2013-07-04 | Method updated | 4000.20.066 1/29/14 | D4951-14 | 65
|  | EN 14538 | 5 point calibration. R>=0.995 | daily | PE Optima 8000 | with use |  |  |  | 2013-07-02 | 2016-07-04 | 4000.20.066 1/29/14 | EN14538 | 66
|  | D5708 | 2 point calibration | daily | PE Optima 8000 | with use |  |  |  | 2013-07-05 | Method updated | 4000.20.066 1/29/14 | D5708-15 | 67
|  | IP 501 | 5 point calibration. R>=0.995 | daily | PE Optima 8000 | with use | [6 empty] | IP 501 | 68
|  | EPA 6010 | 2 point calibration | daily | PE Optima 8000 | with use | [6 empty] | EPA 6010D | 69
|  | UOP 389 | 5 point calibration. R>=0.995 | daily | PE Optima 8000 | with use |  |  |  | 2015-06-07 | 2018-06-07 | 4000.20.099 6/4/15 | UOP 389-14 | 70
|  | MIL | 5 point calibration. R>=0.995, 5-replicates per standard and sample | daily | PE Optima 8000 | with use | [7 empty] | 71
|  | EPA 3050 |  |  | PE Optima 8000 | [7 empty] | EPA 3050B | 72
| Manual Distillation | D86 | Thermometer | every 6 months | [5 empty] | None |  | None | D86-12 | 73
|  |  | Certified reference Material | Every 6 months | [9 empty] | 74
| Vacuum Pump | D6217 | 1-100 kPa less than atmospheric pressure | Annually | S55NXMPF-677 | 2016-05-19 | SR | 2017-05-19 |  | 2015-01-29 | 2018-01-29 | 4000.20.030 2/18/15 | D6217-11 | 75
|  | D7321 | 1-100 kPa less than atmospheric pressure | Annually | S55NXMPF-677 | 2016-05-19 | SR | 2017-05-19 |  | 2012-12-11 | method updated | 4000.20.064 5/23/13 | D7321-14 | 76
|  | D5452 | 1-100 kPa less than atmospheric pressure | Annually | S55NXMPF-677 | 2016-05-19 | SR | 2017-05-19 |  | 2015-01-30 | 2016-01-30 | 4000.20.061 2/2/15 | D5452-12 | 77
|  | D4898 | 550 mm Hg vacuum | Annually | [8 empty] | d4898-16 | 78
|  | FTM 3009 | 14-17 in Hg vacuum | Annually | [9 empty] | 79
|  | D4310 | 13.3 + 0.7 kPa vacuum | annually | [8 empty] | D4310-10(2015) | 80
|  | FTM 3013 | at least 25 in Hg | annually | [9 empty] | 81
| Vacuum Distillation | D1160 (manual) | Calibration of Temp Sensor (annex A1) | annually | [5 empty] | 2013-07-10 | method updated | 4000.20.078 10/21/13 | D1160-15 | 82
|  |  | Temperature response time (annex A2) | annually | [9 empty] | 83
|  |  | Calibration of Vacuum Gage (Annex A3) | annually | [9 empty] | 84
|  |  | n-Hexadecane Check (annex A5) | every 6 months |  | 2016-04-01 | SR | 2016-10-01 | [5 empty] | 85
| Vacuum Distillation | D1160 (automated) | Calibration of Temp Sensor (annex A1) | annually | 6320002 | 2016-04-11 | MCA | 2017-04-11 |  |  |  | Written. Pending review | D1160-15 | 86
|  |  | Temperature response time (annex A2) | annually | 6320002 | [8 empty] | 87
|  |  | Calibration of Vacuum Gage (Annex A3) | annually | 6320002 | [8 empty] | 88
|  |  | n-Hexadecane Check (annex A5) | annually | 6320002 | 2016-10-03 | MC | 2017-10-03 | [5 empty] | 89
| Acidity | D7795B | Flow 10 psi for Nitrogen | daily and verify setting annually with Restek Flow Meter | [5 empty] | Method updated |  | 4000.20.097 2/8/15 | D7795-15 | 90
|  |  | Standardize KOH | monthly |  | 2016-10-26 | JM | 2016-11-26 | [5 empty] | 91
|  | D3242 | Standardize KOH | monthly |  | 2016-10-26 | JM | 2016-11-26 |  | 2014-12-09 | 2015-12-09 | 4000.20.015 1/30/15 | D3242-11 | 92
|  |  | Flow of 600-800 mL/min Nitrogen | daily and verify setting annually with Restek Flow Meter | [9 empty] | 93
| IC | D7328 | See instrument manual |  | 9040317(Dionex) |  |  |  |  | 2013-01-03 | 2016-01-03 | 4000.20.059 3/9/16 | D7328-13 | 94
| IC |  | Calibrate 5 point R >=0.99 | as required by SQC | 9040317(Dionex) | 2016-08-17 | SR | per SQC | [5 empty] | 95
| IC |  | Hot Block 65C | quarterly | [9 empty] | 96
| IC |  | See instrument manual |  | Metrohm | [8 empty] | 97
| IC |  | Calibrate 5 point R >=0.99 | as required by SQC | Metrohm | 2015-11-06 | KAR | per SQC | [5 empty] | 98
| SINDIE | D2622 | multipoint calibration-Gasoline | as required by SQC | BT120321001 | 9/8/2015 | JB | per SQC |  | 2014-06-03 | Method Updated | 4000.20.016 9/16/15 | D2622-16 | 99
|  | D2622 | multipoint calibration-Diesel/Jet | as required by SQC | BT120321001 | 5/12/2016 | SR | per SQC | [5 empty] | 100
|  | D2622 | multipoint calibration-Ethanol | as required by SQC | BT120321001 | 9/8/2015 | JB | per SQC | [5 empty] | 101
|  | D7039 | multipoint calibration-Biodiesel | as required by SQC | BT120321001 | 9/8/2015 | JB | per SQC |  | 2013-08-12 | Method Updated | 4000.20.016 9/16/15 | D7039-15a | 102
|  | D2622/D7039 | Radiation emission | annually | BT120321001 | 9/13/16 | SR | 2017-09-13 | [5 empty] | 103
| SINDIE | D2622 | Radiation emission | annually | BT100430001 (Loaner from XOS) | 7/30/15 | JB | per SQC |  | 2014-06-03 | 2015-06-03 | 4000.20.016 1/31/13 | D2622-10 | 104
|  | D2622 | Radiation emission | annually | BT100430001 (Loaner from XOS) | 7/30/15 | JB | per SQC | [5 empty] | 105
|  | D2622 | Radiation emission | annually | BT100430001 (Loaner from XOS) | 7/31/15 | JB | per SQC | [5 empty] | 106
|  | D7039 | Radiation emission | annually | BT100430001 (Loaner from XOS) | 2015-08-05 | JB | per SQC |  | 2013-08-12 | 2016-08-12 | 4000.20.069 5/30/13 | D7039-13 | 107
| Asoma | D4294 | Radiation emission | annually | 8950 | 9/13/16 | SR | 2017-09-13 | [5 empty] | 108
|  |  | multipoint calibration | As needed per instrument "STD Failed" | 8950 | 10/4/16 | SC | [6 empty] | 109
| Antek 9000 | D5453 | Multipoint Calibration | as required by SQC |  | 2016-10-19 | SC | per SQC | [5 empty] | 110
| Digital Colorimeter | D6045 | Calibrate per Annex X2 | as required by SQC | 98955194 | 2015-09-25 | KAR | per SQC |  | 2015-09-26 | 2016-09-26 | 4000.20.010 9/26/15 | D6045-12 | 111
| Lovibond EC3000 | D6045 | Calibrate per Annex X2 | as required by SQC | 200098 | 2016-07-15 | SC | per SQC | [5 empty] | 112
| Conductivity | D2624 | Calibrate per Annex A2 | Annually | 17?? | 2016-04-01 | SR | 2017-04-01 |  | 2015-09-15 | 2016-09-15 | 4000.20.057 9/15/15 | D2624-15 | 113
|  | D2624 | Calibrate per Annex A2 | Annually | 113242 | 2016-04-01 | SR | 2017-04-01 | [5 empty] | 114
|  | D2624 | Calibrate per Annex A3 | Annually | HI 8633 | 2016-08-01 | JL | 2017-08-01 | [5 empty] | 115
| UV-VIS | D1840 | verify photometric accuracy and wavelength accuracy meets 7.1 | monthly | 2R1R020101 | see chart | KAR | see chart |  | 2015-08-26 | 2016-08-26 | 4000.20.042 5/2/13 | D1840-07(2013) | 116
|  | MIL-PRF-46170 §4.3.3.6 | [11 empty] | 117
|  | D6258 | verify photometric accuracy and wavelength accuracy | annually | [7 empty] | 4000.20.082 | D6258-09(2014) | 118
| MSEP | D3948 | Calibration | Annual | 109803 (Mark X) | 2016-08-09 | EMCEE | 2017-08-09 |  | 2015-02-05 | 2016-02-05 | 4000.20.011 2/1/15 | D3948-14 | 119
| MSEP | D3948 | Calibration | Annual | 15975 (Mark V) | 2016-05-17 | EMCEE | 2017-05-17 |  | 2015-02-05 | 2016-02-05 | 4000.20.011 2/1/15 | D3948-14 | 120
| FIJI | D7797 | CRM verification with one fluid | every six months |  | 2016-08-24 | SR | 2017-02-24 |  | 2014-06-03 | 2017-06-03 | 4000.20.067 5/14/13 | D7797-12 | 121
|  |  | CRM verification with both fluids | annually |  | 2015-10-22 | SR | 2016-10-22 | [5 empty] | 122
|  | D7963 | CRM verification with one fluid | every six months |  | 2016-08-24 | SR | 2017-02-24 |  |  |  |  | D7963-14 | 123
|  |  | CRM verification with both fluids | annually |  | 2015-10-22 | SR | 2016-10-22 | [5 empty] | 124
| Distillation | D86 | Temperature Lag (annex A3) | When new PT100 installed. | (627) 76272614 Unit 1 | 2015-03-17 | JA |  |  | 2015-09-11 | Method updated | 4000.20.002 5/27/13 | D86-15 | 125
|  |  | Hexadecane and Toluene | Every 6 months | (627) 76272614 Unit 1 | 2016-09-14 | NAT/SR | 2017-03-14 | [5 empty] | 126
|  |  | Check Leveler | Every 3 months | (627) 76272614 Unit 1 | 2016-11-03 | SR | 2017-02-03 | [5 empty] | 127
|  |  | Barometer | daily | (627) 76272614 Unit 1 | with use | [7 empty] | 128
|  |  | Temperature Lag (annex A3) | When new PT100 installed. | (627) 97627153 Unit 2 | [8 empty] | 129
|  |  | Hexadecane and Toluene | Every 6 months | (627) 97627153 Unit 2 | 2016-08-05 | JRW | 2017-02-05 | [5 empty] | 130
|  |  | Check Leveler | Every 3 months | (627) 97627153 Unit 2 | 2016-11-03 | MW | 2017-02-03 | [5 empty] | 131
|  |  | Barometer | daily | (627) 97627153 Unit 2 | with use | [7 empty] | 132
|  |  | Temperature Lag (annex A3) | When new PT100 installed. | (627) 98627088 Unit 3 | 2016-10-01 | SR | [6 empty] | 133
|  |  | Hexadecane and Toluene | Every 6 months | (627) 98627088 Unit 3 | 2016-09-28 | NAT | 2017-03-28 | [5 empty] | 134
|  |  | Check Leveler | Every 3 months | (627) 98627088 Unit 3 | 2016-09-28 | NAT | 2016-12-28 | [5 empty] | 135
|  |  | Barometer | daily | (627) 98627088 Unit 3 | with use | [7 empty] | 136
|  |  | Temperature Lag (annex A3) | When new PT100 installed. | (627) 26271510 unit 4 | 2015-02-02 | JA | [6 empty] | 137
|  |  | Hexadecane and Toluene | Every 6 months | (627) 26271510 unit 4 | 2016-06-06 | MW | 2016-12-06 | [5 empty] | 138
|  |  | Check Leveler | Every 3 months | (627) 26271510 unit 4 | 2016-06-13 | SR | 2016-09-16 | This unit is down for repair as of 8/4/2016 |  |  |  |  | 139
|  |  | Barometer | daily | (627) 26271510 unit 4 | with use | [7 empty] | 140
|  |  | Temperature Lag (annex A3) | When new PT100 installed. | (627) 98627105 unit 5 | 2016-09-22 | JRW | [6 empty] | 141
|  |  | Hexadecane and Toluene | Every 6 months | (627) 98627105 unit 5 | 2016-09-22 | JRW | 2017-02-22 | [5 empty] | 142
|  |  | Check Leveler | Every 3 months | (627) 98627105 unit 5 | 2016-09-22 | JRW | 2016-12-22 | [5 empty] | 143
|  |  | Barometer | daily | (627) 98627105 unit 5 | with use | [7 empty] | 144
|  |  | Temperature Lag (annex A3) | When new PT100 installed. | (626) 91626046 | 2015-04-07 | JA |  | Unit down for repair 8/15/15 |  |  |  |  | 145
|  |  | Hexadecane and Toluene | Every 6 months | (626) 91626046 | 2015-08-05 | JA/MJ | 2016-02-05 | Unit down for repair 8/15/15 |  |  |  |  | 146
|  |  | Check Leveler | Every 3 months | (626) 91626046 | 2015-07-07 | SR | 2015-10-07 | Unit down for repair 8/15/15 |  |  |  |  | 147
|  |  | Barometer | daily | (626) 91626008 unit 3 | [8 empty] | 148
|  |  | Thermocouple | Quarterly | (626) 91626008 unit 3 | 2015-02-11 | JA | 2015-05-11 | Note: This unit was taken out of service prior to 5/11/2015 |  |  | 4000.20.006 6/14/11 |  | 149
|  |  | Differential Pressure Cell | N/A | (626) 91626008 unit 3 | N/A | JA | N/A | [5 empty] | 150
|  |  | Metering Pump flow rate (per 11) | daily | (626) 91626008 unit 3 |  |  | with use | [5 empty] | 151
|  |  | Filter Bypass check for leaks | Annually | JFTOT 203 S/N 669 | 2015-03-11 | JA | 2016-03-11 | [5 empty] | 152
| JFTOT | D3241 | Thermocouple | Quarterly | JFTOT 203 S/N 669 | 2016-06-09 | JRW | 2016-09-09 |  | 2015-09-15 | 2016-09-15 | 4000.20.005 1/25/15 | D3241-15e2 | 153
|  |  | Differential Pressure Cell | Annually | JFTOT 203 S/N 669 | 2016-03-09 | JRW | 2017-03-09 | [5 empty] | 154
|  |  | Metering Pump flow rate (per 11) | daily | JFTOT 203 S/N 669 | 2016-03-09 | JRW | with use | [5 empty] | 155
| JFTOT | D3241 | Thermocouple | Quarterly | JFTOT 240 S/N PROTO_8/26/96 | 2016-09-09 | JRW | 2016-12-09 | [5 empty] | 156
|  |  | Differential Pressure Cell | Annually | JFTOT 240 S/N PROTO_8/26/96 | 2016-03-09 | JRW | 2017-03-09 | [5 empty] | 157
|  |  | Metering Pump flow rate (per 11) | daily | JFTOT 240 S/N PROTO_8/26/96 | 2016-03-09 | JRW | with use | [5 empty] | 158
|  |  | Thermocouple | Quarterly | JFTOT 230 S/N 104-9515 | 2016-09-09 | JRW | 2016-12-09 | [5 empty] | 159
|  |  | Differential Pressure Cell | Annually | JFTOT 230 S/N 104-9515 | 2016-03-09 | JRW | 2017-03-09 | [5 empty] | 160
|  |  | Aeration Dryer | Monthly | JFTOT 230 S/N 104-9515 | 2015-04-14 | SR | 2016-04-15 | [5 empty] | 161
|  |  | Differential Pressure Cell | Annually | JFTOT III-1 S/N: 824-0737 | 2016-02-16 | MJ | 2017-02-16 | Down for repair indefinetly |  |  |  |  | 162
|  |  | Metering Pump flow rate (per 11) | daily | JFTOT III-1 S/N: 824-0737 | [8 empty] | 163
|  |  | Thermocouple | Quarterly | JFTOT III-1 S/N: 824-0737 | [8 empty] | 164
| BTU | D240 | Energy equivalent of Calorimeter | Annually | Parr 6200 Calorimeter | 2016-11-11 | NL | 2017-11-11 |  |  |  |  | D240-09 | 165
|  |  | Energy equivalent of calorimeter with volatile fluids | Annually | Parr 6200 Calorimeter | 2016-11-11 | NL | 2017-11-11 | [5 empty] | 166
|  |  | Heat of combustion of pressure tape | Annually | Parr 6200 Calorimeter | 2016-11-10 | NL | 2017-11-10 | [5 empty] | 167
|  |  | Heat of combustion of mineral oil | Annually or with new Lot | Parr 6200 Calorimeter | 2016-08-23 | NL | 2017-08-23 | [5 empty] | 168
| Flash Point | D56 | See Instrument Manual- verify depth of ignitor when replaced. | as needed | HFP 360 S/N 0136E0444 |  |  |  |  | 2015-08-26 | 2016-08-26 | 4000.20.001b 9/11/15 | D56-05(2010) | 169
|  |  | Ensure thermocouple calibration information is updated when PT100 is replaced. | as needed | HFP 360 S/N 0136E0444 | [8 empty] | 170
|  |  | CRM verification | Annually | HFP 360 S/N 0136E0444 | 2016-08-19 | Stanco | 2017-08-19 | [5 empty] | 171
|  |  | Working Standard | Monthly | HFP 360 S/N 0136E0444 | 2016-11-01 | NAT | 2016-12-01 | [5 empty] | 172
| Flash Point Automated | D93 | See Instrument Manual |  | HFP 362 S/N 9836M0042 |  |  |  |  | 2013-02-08 | Method Updated | 4000.20.008 2/12/13 | D93-15 | 173
|  |  | CRM verification | Annually | HFP 362 S/N 9836M0042 | 2016-08-19 | Stanco | 2017-08-19 | [5 empty] | 174
|  |  | Working Standard | Monthly | HFP 362 S/N 9836M0042 | 2016-11-01 | NAT | 2016-12-01 | [5 empty] | 175
|  |  | See Instrument Manual |  | HFP 339 S/N 123390843 | [8 empty] | 176
|  |  | CRM verification | Annually | HFP 339 S/N 123390843 | 2016-03-09 | SR/MS | 2017-03-09 | [5 empty] | 177
|  |  | Working Standard | Monthly | HFP 339 S/N 123390843 | 2016-10-18 | NAT | 2016-11-18 | [5 empty] | 178
| Flash Point Manual | D93A | CRM verification | Annually | HFP 360 S/N 0136E0444 | moved to Mobile Lab |  |  |  | None |  | None |  | 179
|  |  | Working Standard | Monthly | HFP 360 S/N 0136E0444 | [8 empty] | 180
|  |  | verify 3.2-4.8 mm flame | daily | HFP 360 S/N 0136E0444 | [8 empty] | 181
|  |  | Rate of temperature increase 5-6C/min | quarterly | [9 empty] | 182
|  |  | Stirrer Speed | Annually | [9 empty] | 183
|  |  | Working Standard | Monthly |  | 2015-09-01 | SR | 2015-10-01 | [5 empty] | 184
| SetaFlash-Erdco | UN L.2 Sustained Combustability | Thermometer verification at 141 and 167F | annually | 1017 | 2016-09-29 | MCA | 2017-09-29 |  |  |  | Written. Needs review | D3278-96(2011) | 185
|  |  | Timer (1 minute) | Annually | 1017 | 2016-09-29 | MCA | 2017-09-29 | [5 empty] | 186
|  | D3828 | Thermometer verification | annually | 507 | [6 empty] | Written. Needs review | D3828-12a | 187
|  |  | CRM verification | Annually | 507 | [8 empty] | 188
|  |  | Working Standard | Monthly | 507 | [8 empty] | 189
| SetaFlash- Koehler | D3278 | CRM verification | Annually | 507 | 2016-02-14 | AD | 2017-02-14 |  |  |  | Written. Needs review | D3278-96(2011) | 190
|  |  | Thermometer verification | annually | 507 | [8 empty] | 191
|  |  | Working Standard | Monthly | 507 | 2016-11-14 | NL | 2016-12-14 | [5 empty] | 192
|  | D3828 | Thermometer verification | annually | E2005008 | 2016-02-14 | AD | 2017-02-14 |  |  |  | Written. Needs review | D3828-12a | 193
|  |  | CRM verification | Annually | E2005008 | [8 empty] | 194
|  |  | Working Standard | Monthly | E2005008 | 2016-11-14 | NL | 2016-12-14 | [5 empty] | 195
| Grabner FLPH "Naviflash" | D7094 | CRM verification | Monthly | E2005008 | 2015-07-09 | KAR | 2015-08-09 | Taken out of service 7/31/15 |  |  | 4000.20.090 (needs updated to D7094) | D7094-12 | 196
| Automated Cloud/Pour/Freeze Point | D5972, D5949, D5773 | run round robin sample for comparison | Annually | E2005008 | 2016-09-13 | MW | 2017-09-13 |  | 9/15/15 (D5972)- needs completed for D5949 and D5773 | 9/15/16 (D5972) | 4000.20.091 3/23/15 | D5972-05, D5949-14, D5773-10 | 197
| Hach portable Meter | D6756 | Calibrate with multiple standards | annually | E2005008 | 2016-11-09 | SR | 2017-11-09 |  |  |  | 4000.20.043 4/16/15 | D6756-10 (2015) | 198
| Anton Paar DMA 35N | D4052 | Calibrate with Water | quarterly | 07-11-800 | 2014-10-30 | RD | 2015-04-30 | Sent to Floresville | 2015-01-26 | 2016-01-26 | 4000.20.055 2/2/15 | D4052-11 | 199
|  |  | Calibrate with Water | quarterly | 70M3H6C3367C | 2014-10-30 | RD | 2015-04-30 | Taken out of service 12/1/2014 |  |  |  | D4052-11 | 200
| Anton Paar DMA 4500M | D4052 | Calibrate with Water | per SQC | 95020003768 | 2014-11-21 | SR |  |  | 2015-01-26 | 2016-01-26 | 4000.20.055 2/2/15 | D4052-15 | 201
|  | D5002 | Calibration check | weekly | 80573456 (red) | [7 empty] | D5002-16 | 202
| Haze | D4176 | None |  | 587617 (yellow) |  |  |  |  | 2015-03-04 | 2016-03-04 | 4000.20.019 3/4/15 | D4176-04(2014) | 203
| Density by Hydrometer | D1298 | Calibrate Hydrometer | annually | 80865360 |  |  |  |  | 2012-12-05 | 2015-12-05 | 4000.20.012 6/14/11 | D1298-12b | 204
|  |  | Calibrate Thermometer | Annually | [9 empty] | 205
| Density by Pycnometer | D70 | Calibrate Thermometer | Annually | [8 empty] | D70-09e1 | 206
| Doctor | D4952 | None |  | separate sheet |  |  |  |  | 2012-12-05 | 2013-12-05 | 4000.20.017 6/14/11 | D4952-12 | 207
| Liqui-Cult |  | None |  | separate sheet |  |  |  |  | 2013-07-02 | 2016-07-02 | 4000.20.040 |  | 208
| Cold Soak | D7501 | verify 4.5 + 0.5C air bath | daily | separate sheet |  |  |  |  | 2012-12-28 | 2015-12-28 | 4000.20.088 5/22/13 | D7501-12a | 209
|  |  | verify 25 + 0.5C water bath | daily | [9 empty] | 210
| Water by Distillation | D95 | Verify glassware setup prior to first use with water-free solvent distillation | once |  | 2014-04-25 | SR | Moved to Floresville |  |  |  |  | D95-13e1 | 211
| Sediment by Extraction | D473 | None |  |  | moved to Floresville |  |  |  | 2014-06-05 | 2017-06-05 |  | D473-07(2012) | 212
| Cetane Index | D976 | [10 empty] | D976-06(2011) | 213
| Cetane Index | D4737 |  |  | N/A | [7 empty] | D4737-10 | 214
| BTU- Estimated | D3338 | [10 empty] | D3338/D3338M-09(2014)e1 | 215
| RVP | D323 | [10 empty] | D323-08(2014) | 216
| FSII | D5006 | [10 empty] | D5006-11 | 217
| Saybolt Color | D156 | Color Standards Discs | upon purchase of standards |  |  | EA |  |  | 2015-09-10 | 2016-09-10 | 4000.20.086 9/10/15 | D156-15 | 218
| ASTM Color | D1500 | Color Standard Disc | upon purchase of standards | [5 empty] | 2014-12-23 | 2017-12-23 | 4000.20.087 9/11/14 | D1500-12 | 219
| Automated Freeze Point | D5972 | cross-check with ASTM round robin sample | annual |  | 2014-10-13 | KAR | moved to Floresville |  | 2015-02-05 | 2016-02-05 | 4000.20.091 3/23/15 | D5972-05(10) | 220
| Kauri Butanol | D1133 |  | daily with use | N/A |  | SB | [5 empty] | D1133-13 | 221
| Mini-Flash(Naviflash) | D6450 | Analyze standard of Anisole or Dodecane | daily with use | N/A | moved to Floresville |  |  |  | 2014-09-28 | 2017-09-28 | 4000.20.090 9/28/14 | D6450-12 | 222
| Calorimeter (BTU) | D240 | Standardize NaOH | monthly | 70N346C3908C | 2016-11-07 | NL | 2016-12-07 |  |  |  | 4000.20.105 12/3/15 | D240-14 | 223
|  | D808 | none outside of SQC | [8 empty] | 4000.20.107 3/7/16 | D808-11 | 224
|  | D665 | temperature of bath 60 + 1C | annually | 07-11-339 | [6 empty] | 4000.20.072 7/21/13 | D665-14e1 | 225
|  |  | Speed of chuck on grinder 1700-1800 rpm | annually |  | 2016-04-15 | SR | 2017-04-15 | [5 empty] | 226
|  |  | Stirrer apparatus speed 1000 + 50 rpm | annually |  | 2016-04-15 | SR | 2017-04-15 | [5 empty] | 227
| Shaker Bath #1 | D665 |  | Annually | Shaker Bath | 2016-03-01 | SR | 2017-03-01 |  |  |  | 4000.20.072 (out of date) | D665-14e1 | 228
| Autotitrator | D664A/B | 10mL burette with + 0.02mL accuracy | quarterly | Titrando 809 SN10809001013164 | 2015-08-05 | SA | 2015-11-05 |  | 2015-11-02 | 2016-11-02 | 4000.20.068 7/29/13 | D664-11a | 229
|  |  | 10mL burette with + 0.02mL accuracy | quarterly | Titrando 905 | [8 empty] | 230
|  | D664A/B | Electrodes-mV difference between 4 and 7 buffer >= 162 | daily with calibration | [9 empty] | 231
|  | D664B | Standardize KOH | weekly | Titrando 809 SN10809001013164 | 2016-10-28 | NAT | 2016-11-04 | [5 empty] | 232
|  | D664A | Standardize KOH | bi-weekly | Titrando 905 | 2016-11-18 | NL | 2016-12-02 | [5 empty] | 233
|  | D664A/B | Standardize with pH 4, 7 and 11 | daily | [9 empty] | 234

## Middle Lab
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| GC-FID | D4815 | Calibration | on setup and as required by SQC | 2921A23364 | 2015-03-07 | SR | as required by SQC |  | 2013-06-27 | Method updated | 4000.20.050 5/27/13 | D4815-15b | 1
| GC-TCD | D3606 | Calibration | on setup and as required by SQC | 3336A54446 | 2016-05-24 | SR | as required by SQC |  | 2012-12-01 | 2015-12-01 | 4000.20.051 11/18/13 | D3606-10 | 2
| GC-MS | D5769 | Calibration | on setup and as required by SQC | 3336A55329 |  | SR | as required by SQC |  | 2013-06-27 | Method updated | 4000.20.052 5/27/13 | D5769-15 | 3
|  |  | Tune MS | prior to calibration |  |  | SR | as required by SQC | [5 empty] | 4
| Viscosity | D445 | Viscometer Verification | annually | See separate sheet |  |  |  |  | 2015-09-15 | 2016-09-15 | 4000.20.024 9/15/15 | D445-15a | 5
|  | D445 | Temperature of Bath between 15-100C, + 0.02 C. For all other temps, + 0.05C | annually | Box on high temperature bath | 2016-05-26 | MS | 2017-05-26 | [5 empty] | 6
[5 empty] | Box on 40C bath | 2016-02-02 | SC | 2017-02-02 | [5 empty] | 7
[5 empty] | Jet bath | 2016-02-02 | SR | 2017-02-02 | [5 empty] | 8
[5 empty] | Box on Ultra Low Temp Bath | 2016-05-25 | MS | 2017-05-25 | [5 empty] | 9
[5 empty] | Box on Igloo Cooler Bath | 2016-02-02 | SR | 2017-02-02 | [5 empty] | 10
[5 empty] | Box on 100C bath | 2016-08-11 | JRW | 2017-08-11 | [5 empty] | 11
| Koehler | D445 |  |  | R09800233 | 2016-08-22 | JL | 2017-08-22 | [5 empty] | 12
|  |  | Timer + 0.1s | annually | See separate sheet | [8 empty] | 13
|  |  | Viscosity of Silicone Oil <200 cSt @ 100C | Quarterly | high temp bath | [8 empty] | 14
| Viscosity Stability | D2532 | All items for D445 | [9 empty] | D2532-14 | 15
|  |  | Verify bath temp + 0.03 C of reported temperature for entire run | daily | [9 empty] | 16
| Drying Oven | D6217 / D2276/ D5452/ D7321 | Temperature 90 + 5 C | daily | 2441 |  | JA | [6 empty] | 17
|  | D4898 | Temperature 80 C | daily | [9 empty] | 18
|  | D4310 | Temperature 105C | daily | [9 empty] | 19
|  | D2274 | Temperature 105 + 5 C | daily | [8 empty] | D2274-14 | 20
|  | D2274 | Temperature 80 + 2 C | daily | [9 empty] | 21
|  | D2440 | Temperature 105-110C | daily | [8 empty] | D2440-13 | 22
|  | FTM 3010/ 3012 | Temperature of 90C | daily | [8 empty] | FTM 3010.1/ 3012 | 23
| Cloud Point | D5771 | Temperature Calibration | Annually | 812 | 2014-09-10 | SR | 2015-09-10 | removed from service 1/26/15 |  |  |  | D5771-12 | 24
| Water Sludging | MIL-DTL-17111E 4.8.3.5 | Requirements for D445 | daily | [5 empty] | 2015-03-25 | 2018-03-25 | 4000.20.102 6/25/15 | MIL-DTL-17111E 4.8.3.5 | 25
|  |  | Paddle rotation at 5000 + 200 rpm | annually | [9 empty] | 26
|  |  | Bath at 38C | daily | [9 empty] | 27
| Automatic Viscometer | D445 | Temperature 40 + 0.2 C | Annually | 74810395 | 2014-10-23 | KAR | 2015-10-23 | removed from service 1/26/15 |  |  |  |  | 28
|  | D445 | Timing | Annually |  | 2014-10-23 | KAR | 2015-10-23 | [5 empty] | 29
| HLPS- Rig 1 | SAE ARP 5996C | Flow rate 1mL/min | daily | [9 empty] | 30
|  |  | Calibrate Thermocouple with Lead and Tin | every 6 months |  | 2016-10-17 | JB/JRW | 2017-04-17 |  |  |  | 4000.20.037 4/15/15 | SAE ARP 5996C | 31
|  |  | Oven at 100 +/- 5 C | daily | [9 empty] | 32
| Balance-Radwag |  |  | semi-annual |  | 2015-03-18 | MS | 2015-09-18 | Returned to manufacturer |  |  |  |  | 33
|  |  | Drive pump flow rate | every 6 months | [9 empty] | 34
| HLPS- Rig 2 | SAE ARP 5996C | Flow rate 1mL/min | daily | 030-0517 | [8 empty] | 35
|  |  | Calibrate Thermocouple with Lead and Tin | every 6 months |  | 2015-04-17 | BB | 2015-10-17 |  |  |  | 4000.20.037 4/15/15 | SAE ARP 5996C | 36
|  |  | Oven at 100 +/- 5 C | daily | [9 empty] | 37
|  |  | Drive pump flow rate | every 6 months | [9 empty] | 38
| HLPS- Rig 3 | SAE ARP 5996C | Flow rate 1mL/min | daily | 028-0407 | [8 empty] | 39
|  |  | Calibrate Thermocouple with Lead and Tin | every 6 months |  | 2016-10-17 | JB/JRW | 2017-04-17 | [5 empty] | 40
|  |  | Oven at 100 +/- 5 C | daily | [9 empty] | 41
|  |  | Drive pump flow rate | every 6 months | [9 empty] | 42
| Particle Count | ISO 4006 NAS 1638 |  |  |  | 2014-12-23 | BB | [6 empty] | 43
| Balance- Vibra |  |  | every 6 months | 138975036 | 2016-02-03 | Aldinger | 2016-08-03 | [5 empty] | 44
| Igloo Cooler Bath |  |  | Annually |  | 2016-10-15 | JRW | 2017-10-15 | [5 empty] | 45
| Balance-Tree |  |  | every 6 months | HR1409504 | 2016-02-03 | Aldinger | 2016-08-03 | [5 empty] | 46
| Micrometer | [12 empty] | 47
| Copper Corrosion Bath | D130 | 50 +/- 1 C | daily | [8 empty] | D130-12 | 48
|  |  | copper strip dimensions (10.5-14.5mm width, 1.5-3.2mm thick, 70-80mm length) | Quarterly |  | 2016-09-08 | JB | 2016-12-08 | [5 empty] | 49
|  |  | Check ASTM Copper Strip Corrosion Standard Chart for Fading | annually |  | 2016-09-08 | JB | 2016-12-08 | [5 empty] | 50
|  |  | Check Copper Strip for scratches, pits, etc. | daily | [9 empty] | 51
| TGA | E1269 |  | every 6 months | 10100682000014 | 2016-10-27 | MCA | 2017-04-27 |  |  |  |  | E1269-11 | 52
|  |  | Calibrate Temp. Corr. | every 6 months | 10100682000014 | 2016-10-12 | MCA | 2017-04-12 | [5 empty] | 53
| Foaming | D892 | [9 empty] | Written. Needs review | D892-13 | 54
| Sonicator |  |  |  | Branson | [8 empty] | 55
|  |  | Glassware Oven 150C | quarterly | QV1-20761 | 2016-09-28 | SR | 2016-12-28 | [5 empty] | 56
| Humidity Cabinet | D1748 | Relative humidity 50% | daily | [7 empty] | 4000.20.080 (out of date) | D1748-10 (2015) | 57
| Humidity Cabinet | D1748 | Temperature 27C | daily | [9 empty] | 58
| Humidity Cabinet | D1748 | rpm | annually |  | 2016-11-07 | RL | 2017-11-07 | [5 empty] | 59
| Humidity Cabinet | D1748 | Temperature 49C | annually |  | 2016-11-07 | RL | 2017-11-07 | [5 empty] | 60
| Instron Tensile Tester | D412 | rate of grip separation 500 ± 50 mm/min | annually | 193 | 2016-02-03 | Aldinger | 2017-02-03 |  |  |  |  | ASTM D412-15a (2013) | 61
|  |  | Dynamometer to measure force within ± 2% | annually | 193 | 2016-02-03 | Aldinger | 2017-02-03 | [5 empty] | 62
| Isoteniscope | D2879 | Temperature +/-1.5C | annually |  | 2016-06-22 | MS | 2017-06-22 | [5 empty] | 63
| Isoteniscope | D2879 | Pressurte +/- 1torr | annually |  | 2016-06-22 | MS | 2017-06-22 | [5 empty] | 64

## Middle Hot Room
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
Hot Room | Muffle Furnace | D482 | Temperature 775 + 25 C | Annually | 62700 | 2016-05-19 | MS | 2017-05-19 |  | 2012-12-07 |  | 4000.20.020 1/3/13 | D482-13 | 1
|  | D874 | Temperature 775 + 25 C | Annually | 62700 | 2016-05-19 | MS | 2017-05-19 |  |  |  | 4000.20.096 | D874-13a | 2
|  | IP 501 | Temperature 525 + 25C | Annually | [9 empty] | 3
|  |  | Temperature 925 + 25C | Annually | [9 empty] | 4
|  | FTM 3010.1 | Temperature 810C | Annually | [8 empty] | FTM 3010.1 | 5
Hot Room | Muffle Furnace | D482 | Temperature 775 + 25 C | Annually | BA-1386 | 2014-04-03 | MS | 2015-04-03 | This instrument was damaged before 4/3 and is awaiting repairs |  |  |  |  | 6
|  | D874 | Temperature 775 + 25 C | Annually | BA-1386 | 2014-04-03 | MS | 2015-04-03 | This instrument was damaged before 4/3 and is awaiting repairs |  |  |  |  | 7
|  | IP 501 | Temperature 525 + 25C | Annually | [9 empty] | 8
|  |  | Temperature 925 + 25C | Annually | [9 empty] | 9
|  | FTM 3010.1 | Temperature 810C | Annually | [9 empty] | 10
| MCRT | D4530 | Temperature ramp 10-40C/min | annually | 00068 | 2016-03-11 | SR | 2017-03-11 |  | 2012-12-12 | 2015-12-12 | 4000.20.025 7/29/13 | D4530-11 | 11
|  |  | Nitrogen flow rate 600mL/min and 150 mL/min | annually | 00068 | 2016-03-11 | SR | 2017-03-11 | [5 empty] | 12
|  |  | Temperature 500C | annually | 00068 | 2016-03-11 | SR | 2017-03-11 | [5 empty] | 13
|  |  | Temperature ramp 10-40C/min | annually |  |  |  |  | Unit down for repair 9/1/2015 |  |  |  |  | 14
|  |  | Nitrogen flow rate 600mL/min and 150 mL/min | annually | 649 | 2015-05-22 | MS | 2016-05-22 | Unit down for repair 9/1/2015 |  |  |  |  | 15
|  |  | Temperature 500C | annually | 649 | 2015-05-22 | MS | 2016-05-22 | Unit down for repair 9/1/2015 |  |  |  |  | 16
| C&O | D4636- Procedure 1 | Air Flow 10 L/h | Quarterly | Yellow Block (old) |  |  |  |  | 2015-03-01 | 2018-03-01 | 4000.20.071 2/19/15 | D4636-14 | 17
|  | D4636-Procedure 2 | Air Flow 5 L/h | Quarterly | Yellow Block (old) | [8 empty] | 18
|  | D4636 | Air Flow 10 L/h | Quarterly | Yellow Block (old) | [8 empty] | 19
|  | D4636 | Temperature + 0.5C | per test | Yellow Block (old) | 2015-05-29 | JW | 2016-05-29 | [5 empty] | 20
|  | D4636 | Condenser water temp 18 + 3 C | quarterly | Yellow Block (old) | [8 empty] | 21
| Oil Bath | D4310 | Temperature 95 + 0.2 C | annually |  | 2016-11-15 | SC | 2017-11-15 |  |  |  |  | D4310-10 (2015) | 22
|  | D4310 | Oxygen Flow 3 + 0.1 L/h | annually |  | 2016-11-15 | SC | 2017-11-15 | [5 empty] | 23
|  | MIL-PRF-14107D | Temperature 100 + 0.5 C | annually |  | 2016-11-15 | SC | 2017-11-15 | [5 empty] | 24
|  | MIL-PRF-14107D | air flow 5 + 0.2 L/hr | annually |  | 2016-11-15 | SC | 2017-11-15 | [5 empty] | 25
|  | MIL-DTL-17111D | Temperature 93C | annually |  | 2016-11-15 | SC | 2017-11-15 | [5 empty] | 26
|  | MIL-DTL-17111D | dry air 12 + 1 g/hr | annually |  | 2016-11-15 | SC | 2017-11-15 | [5 empty] | 27
|  | D2274 | Temperature 95 + 0.2 C | daily | [8 empty] | D2274-14 | 28
|  |  | O2 Flow 3 + 0.3 L/h | daily | [9 empty] | 29
|  | D7462 | Temperature 95 + 0.2 C | daily | [8 empty] | D7462-11 | 30
|  |  | O2 Flow 3 + 0.1 L/h | daily | [9 empty] | 31
| X-proof oven | FTM 5305 | Temperature 232 + 3 C | annually | CSP-67 | 2016-11-11 | JRW | 2017-11-11 |  |  |  |  | FTM 5305.1 | 32
|  | ASTM D471 | Temperature as requested + 2C | annually | [8 empty] | ASTM D471-12a | 33
|  | FTM 3403 | Temperature 105 + 1 C | annually | [8 empty] | FTM 3403.2 | 34
| SOD | FTM 5321 | Temperature 163 + 1C | annually | BE11 | 2016-08-11 | JW | 2017-08-11 |  |  |  | 4000.20.074 7/21/13 |  | 35
|  | FTM 5321 | Motor 600 + 50 rpm | annually |  | 2016-11-09 | JL | 2017-11-09 | [5 empty] | 36
|  | FTM 5321 | Air Flow 0.94 + 0.047 L/min | annually | 401395-2 | 2016-08-11 | JL | 2017-08-11 | [5 empty] | 37
| Oxidation Bath | D943 | Temperature 95 + 0.2 C | annually | [9 empty] | 38
|  |  | O2 Flow 3 + 0.1 L/h | annually | [9 empty] | 39
| Cleveland Open Cup Flash/Fire Point | D92 | temperature measuring device 6.4 + 0.1 mm from bottom of cup | daily |  |  |  |  | Instrument was taken out of service 10/21/15 |  |  | 4000.20.095 2/5/15 | D92-12b | 40
|  |  | Certified Reference Material | annually | 2015-01-X3 | 2015-03-02 | JL | 2016-03-02 | Instrument was taken out of service 10/21/15 |  |  |  |  | 41
|  |  | Secondary Working Standard | monthly |  | 2015-09-30 | JL | 2015-10-30 | Instrument was taken out of service 10/21/15 |  |  |  |  | 42
|  |  | flame size 3.2-4.8 mm | daily |  |  |  |  | Instrument was taken out of service 10/21/15 |  |  |  |  | 43
|  |  | temperature rise until 100C below flash, 5-17 C/min | quarterly |  | 2015-09-30 | JL | 2015-12-30 | Instrument was taken out of service 10/21/15 |  |  |  |  | 44
|  |  | temperature rise above 100C below flash, 5-6 C/min | quarterly |  | 2015-09-30 | JL | 2015-12-30 | Instrument was taken out of service 10/21/15 |  |  |  |  | 45
| Cleveland Open Cup Flash/Fire Point | D92 | temperature measuring device 6.4 + 0.1 mm from bottom of cup | daily | Stanhope Seta Multi-flash 343000-0 | [6 empty] | 4000.20.095 2/5/15- written for old unit | D92-12b | 46
|  |  | Certified Reference Material | annually | 2015-01-X3 | 2016-04-20 | SR | 2017-04-20 | [5 empty] | 47
|  |  | Secondary Working Standard | monthly |  | 2016-11-02 | JL | 2016-12-02 | [5 empty] | 48
|  |  | flame size 3.2-4.8 mm | daily | Stanhope Seta Multi-flash 343000-0 | [8 empty] | 49
|  |  | temperature rise until 100C below flash, 5-17 C/min | quarterly | Stanhope Seta Multi-flash 343000-0 | 2016-11-02 | JL | 2017-02-02 | [5 empty] | 50
|  |  | temperature rise above 100C below flash, 5-6 C/min | quarterly | Stanhope Seta Multi-flash 343000-0 | 2016-11-02 | JL | 2017-02-02 | [5 empty] | 51
| Cleveland Open Cup Flash/Fire Point | D92 | temperature measuring device 6.4 + 0.1 mm from bottom of cup | daily | Manual | [6 empty] | 4000.20.095 2/5/15- written for old automated unit | D92-12b | 52
|  |  | Certified Reference Material | annually | Manual | 2016-11-10 | MC | 2017-11-10 | [5 empty] | 53
|  |  | Secondary Working Standard | monthly | Manual | 2016-11-10 | MC | 2017-11-10 | [5 empty] | 54
|  |  | flame size 3.2-4.8 mm | daily with use | Manual | [8 empty] | 55
|  |  | temperature rise until 100C below flash, 5-17 C/min | daily with use | Manual | [8 empty] | 56
|  |  | temperature rise above 100C below flash, 5-6 C/min | daily with use | Manual | [8 empty] | 57
| Copper Corrosion | D130 | Bath temp 100C | Annually |  | 2016-02-02 | SR | 2017-02-02 |  | 2015-02-17 | 2016-02-17 | 4000.20.003 2/24/15 | D130-12 | 58
|  | D130 | Dimensions of Copper Strips | Annually |  | 2016-09-07 | RL | 2017-09-07 | [5 empty] | 59
| Induction | D525 | 100C-Temp | Annually | [7 empty] | 4000.20.031 | D525-12a | 60
|  |  | Pressure gauge | Annually | 21816070068 | 2016-09-12 | Additel | 2017-09-12 | [5 empty] | 61
|  |  | Pressure gauge | Annually | 218160C0038 | 2016-09-12 | Additel | 2017-09-12 | [5 empty] | 62
|  |  | Pressure gauge | Annually | 218160C0010 | 2016-09-12 | Additel | 2017-09-12 | [5 empty] | 63
|  |  | Pressure gauge | Annually | 218160C0019 | 2016-09-12 | Additel | 2017-09-12 | [5 empty] | 64
|  |  | Verify Timing for recorder | Annually | Recorder 1 | [8 empty] | 65
|  |  | Verify Timing for recorder | Annually | Recorder 2 | [8 empty] | 66
| Elastomer Compatibility | FTM 3604 | Oven at Temperature requested + 1C | with use | [5 empty] | 2015-02-05 | 2018-02-05 | 4000.20.076 2/4/15 | FTM 3604.2 | 67
| RPVOT Digital Pressure Gauge | D2272 | ±1 PSI annually | Quarterly | 16213493 | 2016-09-12 | Dickson | 2016-12-12 | [5 empty] | 68
| RPVOT | D2272 | Rotation 100 ±5 rpm | Annually |  | 2016-09-20 | AC | 2017-09-20 | [5 empty] | 69
| RPVOT | D2272 | Bath temp 150 ± 6C | semi-annually |  | 2016-09-20 | AC | 2017-03-20 | [5 empty] | 70
| Grease Ox-Stab | D942 | Bath Temp 99C | annually |  | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 71
| Grease Ox-Stab | D942 | Pressure gauge | annually | [9 empty] | 72
| Grease Ox-Stab | D942 | Pressure gauge | annually | [9 empty] | 73
| Grease Ox-Stab | D942 | Pressure gauge | annually | [9 empty] | 74
| Grease Ox-Stab | D942 | Pressure gauge | annually | [9 empty] | 75
| Ramsbottom Furnace | D524 | Confirm Heating Singly and fully loaded | annually | N/A | 11/7/2016 | MS | 2017-11-07 | [5 empty] | 76
|  | D524 | Confirm thermocouple | every 50 hours of run time | [9 empty] | 77

## Back Hot Room
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| C&O | D4636- Procedure 1 | Air Flow 10 L/h | Quarterly | 6-place block | 2016-09-06 | JRW | 2016-12-06 |  | 2015-03-01 | 2016-03-01 | 4000.20.071 2/19/15 | D4636-14 | 1
|  | D4636-Procedure 2 | Air Flow 5 L/h | Quarterly | 6-place block | 2016-09-06 | JRW | 2016-12-06 | [5 empty] | 2
|  | D4636 | Air Flow 10 L/h | Quarterly | 6-place block | 2016-09-06 | JRW | 2016-12-06 | [5 empty] | 3
|  | D4636 | Temperature + 0.5C | per test | 6-place block | [8 empty] | 4
|  | D4636 | Condenser water temp 18 + 3 C | quarterly | 6-place block | 2016-11-03 | JL | 2017-02-03 | [5 empty] | 5
| C&O | D4636- Procedure 1 | Air Flow 10 L/h | Quarterly | 12-place block | 2016-09-14 | JW | 2016-12-14 | [5 empty] | 6
|  | D4636-Procedure 2 | Air Flow 5 L/h | Quarterly | 12-place block | 2016-09-14 | JW | 2016-12-14 | [5 empty] | 7
|  | D4636 | Air Flow 10 L/h | Quarterly | 12-place block | 2016-09-14 | JW | 2016-12-14 | [5 empty] | 8
|  | D4636 | Temperature + 0.5C | per test | 12-place block | 2016-03-11 | JRW | 2017-03-11 | [5 empty] | 9
|  | D4636 | Condenser water temp 18 + 3 C | quarterly | 12-place block | 2016-11-03 | JL | 2017-02-03 | [5 empty] | 10
| IP 280/D2440 block | D2440 | Temperature 110 + 0.5 C | annually | NA | 2016-03-22 | MS | 2017-03-22 |  |  |  |  | IP 280, | 11
|  | D2440 | Oxygen Flow 1.0 + 0.1 L/h | annually |  | 2016-03-22 | MS | 2017-03-22 | [5 empty] | 12
|  | IP 280 | Temperature 110 + 0.5 C | annually |  | 2016-03-22 | MS | 2017-03-22 | [5 empty] | 13
|  | IP 280 | Oxygen Flow 1.0 + 0.1 L/h | annually |  | 2016-03-22 | MS | 2017-03-22 | [5 empty] | 14
| Thermal Stability | FTM 2508 | Block Temperature 260 + 1C | annually | [9 empty] | 15
|  | FTM 3411 | Block Temperature 274 + 1C | annually | NA | 2016-02-05 | MS | 2017-02-05 |  |  |  | 4000.20.048 10/11/11 |  | 16
|  | FTM 3411 | Vacuum pump less than 10um mercury | annually | [9 empty] | 17
| Evaporation | D972 | Air Flow 2.58 + 0.02 g/min (2L/min) | annually |  | 2016-07-20 | JW | 2017-07-20 |  |  |  | 4000.20.045 8/4/11 | D972-02 (2008) | 18
| Evaporation | D972 | Temperature | annually |  | 2016-07-20 | JW | 2017-07-20 | [5 empty] | 19
| Elastomer Block (1) | FTM 3432, 3433 | Temperature | annually | Front one of block with two elastomers | 2016-09-14 | JRW | 2017-09-14 |  |  |  | 4000.20.044 out of date | FTM 3432.1, 3433 | 20
| Elastomer Block (2) | FTM 3432, 3433 | Temperature | annually | Back one of block with two elastomers | 2016-09-14 | JRW | 2017-09-14 |  |  |  |  | FTM 3432.1, 3433 | 21
| Elastomer Block (3) | FTM 3432, 3433 | Temperature | annually | In block with TSC | 2016-09-14 | JRW | 2017-09-14 |  |  |  |  | FTM 3432.1, 3433 | 22
| high temp stability | MIL-PRF- 87252C | Temperature 200 +/- 2 C | daily | [7 empty] | Under review |  | 23
|  |  | 1L /hr nitrogen flow | annually | [9 empty] | 24
| Shaker Bath #2 | FTM 3411 |  | Annually | 601051369 | 2016-07-18 | SR | 2017-07-18 | [5 empty] | 25
| JFTOT 215 |  | Metering Pump flow rate (per 11) | daily | JFTOT 215 S/N 727 | 2015-03-16 | JA | with use | [5 empty] | 26
|  |  | Thermocouple | Quarterly |  | 2016-11-14 | JRW | 2017-02-14 | [5 empty] | 27
|  |  | Differential Pressure Cell | Annually |  | 2016-11-14 | JRW | 2017-11-14 | [5 empty] | 28
| HLPS- Crude |  | Calibrate Thermocouple with Lead and Tin | every 6 months |  | 2016-11-11 | JRW | 2017-05-11 | [5 empty] | 29
|  |  | Oven at 100 +/- 5 C | daily | [9 empty] | 30
|  |  | Drive pump flow rate | every 6 months |  | 2016-11-11 | JRW | 2017-05-11 | [5 empty] | 31

## Back Lab
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| Large Oven 1 | FTM 3604 | Temperature +/- 1C | annually | 010K-473514-0K | 2016-11-14 | JRW | 2017-11-14 |  |  |  |  | FTM 3604.2 | 1
| Large Oven 2 | FTM 3604 | Temperature +/- 1C | annually | [9 empty] | 2
| Large Oven 3 | FTM 3604 | Temperature +/- 1C | annually | z05K-507671-ZK | 2016-11-14 | JRW | 2017-11-14 | [5 empty] | 3
| Large Oven 4 | FTM 3604 | Temperature +/- 1C | annually | Y28K-505957-ZK | 2016-11-11 | JRW | 2017-11-11 | [5 empty] | 4
| Large Oven 5 | FTM 3604 | Temperature +/- 1C | annually | 704o0054 | 2016-11-11 | JRW | 2017-11-11 | [5 empty] | 5
| Medium Oven 1 (Oven 10) |  | Temperature +/- 1C | annually | V0IT-520429-VT | 2016-11-11 | JW | 2017-11-11 | [5 empty] | 6
| Medium Oven 2 |  | Temperature +/- 1C | annually | V0IT-520430-VT | 2016-08-25 | JL | 2017-08-25 | [5 empty] | 7
| Rotiserie Oven |  | Temperature +/- 1C | annually | Labline 3479-1 SN:0010 | 2016-10-17 | JRW | 2017-10-17 | [5 empty] | 8
| Humidity Chamber/Air Bath | D7501 | 4.5 C +/- 0.5C | annually |  | 2016-04-22 | MS | 2017-04-22 |  |  |  | 4000.20.088 | D7501-12a | 9
| Autoignition | E659 | Temperature up to 350 C +/- 1 C, Temperature above 350 C +/- 2c | Daily with use | 889181 | [6 empty] | 4000.20.101 | E659-15 | 10
| Sonic Shear | D2603 |  |  | WW301A00029 | [7 empty] | D2603-01 (2013) | 11
| Sonic Shear | D2603 |  |  | 11170 | [7 empty] | D2603-01 (2013) | 12
| Salt Spray | B117 | Air pressure between 10-25 psi | Annually | [7 empty] | Under review | B117-11 | 13
|  |  | Temperature 35 + 1.1-1.7C | twice daily at least 7 hours apart | [9 empty] | 14
|  |  | 1-2ml atomized per hour per 80 sq cm over 16 hours | daily | [9 empty] | 15
| Flame Propagation | D5306 | [10 empty] | D5306-92 (2013) | 16
| High Temperature, High Pressure Spray Ignition | FTM 6052 | [11 empty] | 17
| Wick Flammability | Boeing Procedure 8.7.1 | [9 empty] | Under review |  | 18
| Manifold Ignition |  | Maintain Temperature of 704.4 ± 17C | annually | N/A | 2016-05-30 | MS | 2017-05-30 | [5 empty] | 19
| Isoteniscope | D2879 | Calibrate air bath over ambient to 748K +/- 2K | annually | Unit 1 | [8 empty] | 20
|  |  | Pressure Transducer |  | Unit 1 | [8 empty] | 21
|  |  | Thermocouple and reader | annually | Unit 1 | [8 empty] | 22
| Thermal Conductivity | [12 empty] | 23
| 4-Ball | D4172/ D2266 | spindle speed 1200 + 60 rpm | Annual | 184-682-VD | 2016-05-13 | MS | 2017-05-13 |  | 2011-05-20 | 2014-05-20 | 400.20.046 9/20/11 | D4172-94(2016) | 24
|  |  | temperature 75 + 2 C | Annual | 184-682-VD | 2016-05-13 | JL/MS | 2017-05-13 | [5 empty] | 25
|  |  | timer 60 + 1 minute | Annual | 184-682-VD | 2016-05-13 | JL | 2017-05-13 | [5 empty] | 26
|  |  | balanced loading mechanism | Annual | 184-682-VD | 2016-05-13 | MS | 2017-05-13 | [5 empty] | 27
| BOCLE | D5001 | Temperature 25 + 1C | Annual | 55 | 2016-04-06 | SR | 2017-04-06 |  | 2013-08-30 | 2016-08-30 | 4000.20.075 9/9/13 | D5001-10(2014) | 28
|  |  | Relative Humidity 10 +0.2% at 25C | Annual | 55 | 2016-04-06 | SR | 2017-04-06 | [5 empty] | 29
|  |  | cylinder 240 + 1 rpm | Annual | 55 | 2016-04-06 | SR | 2017-04-06 | [5 empty] | 30
|  |  | timer 30 + 0.1 min | Annual | 55 | 2016-04-06 | SR | 2017-04-06 | [5 empty] | 31
|  |  | Reference Fluids (see method for spec) | each new batch | [9 empty] | 32
|  |  | Test ring (see method for spec) | each new ring | [9 empty] | 33
|  |  | Air Flow 0.5 L/min flow through, 3.3 L/min over the fluid | Annual | 55 | 2016-04-06 | SR | 2017-04-06 | [5 empty] | 34
| Microscope | D4172/ D2266/ D5001/ D6079 | accuracy 0.01 mm | Annual | 10446339 | 2016-12-05 | MS | 2017-12-05 | [5 empty] | 35
| HFRR | D6079 | Stroke length 1 + 0.02 mm | Quarterly | D506 | 2016-08-23 | SC | 2016-11-23 |  |  |  | 4000.20.073 7/21/13 | D6079-11 | 36
|  |  | frequency 50 + 1 Hz | Quarterly | D506 | 2016-08-23 | SC | 2016-11-23 | [5 empty] | 37
|  |  | temp 60 + 2 C | Annually | D506 | 2016-02-24 | SC | 2017-02-24 | [5 empty] | 38
|  |  | timer 75 + 1 min | Annually | D506 | 2016-02-24 | SC | 2017-02-24 | [5 empty] | 39
|  |  | Fluid A & B | every 20 samples | [9 empty] | 40
| FIA | D1319 | Gas pressure 0-103 kPa | daily | [5 empty] | 2015-08-20 | ASAP- method updated | 4000.20.033 9/10/15 | D1319-15 | 41
| Centrifuge | D2709 | rcf of 800 + 60 | Annual | Koehler K60002 | 2016-03-11 | SR | 2017-03-11 |  | 2014-06-02 | 2017-06-02 | 4000.20.035 6/15/11 | D2709-96(2011)e1 | 42
|  | D1796 | rcf 500-800 | Annual | Koehler K60002 | 2016-03-11 | SR | 2017-03-11 |  | 2015-02-10 | 2018-02-10 | 4000.20.098 | D1796-11e1 | 43
|  |  | temperature 60 + 1C | Annual | Koehler K60002 | 2016-03-11 | SR | 2017-03-11 | [5 empty] | 44
|  | D4007 | rcf of 600 minimum | Annual | Koehler K60002 | 2016-03-11 | SR | 2017-03-11 |  |  |  | 4000.20.100 | D4007-11e1 | 45
|  |  | temperature 60 + 1C | Annual | Koehler K60002 | 2016-03-11 | SR | 2017-03-11 | [5 empty] | 46
|  | D4636 | rcf 840 + 40 | Annual |  | 2016-03-11 | SR | 2017-03-11 |  |  |  | 4000.20.071 7/21/13 | d4636-14 | 47
|  | MIL-DTL- 17111E§4.5.3.4 | 1200 rpm | Annual |  | 2016-03-11 | SR | 2017-03-11 |  |  |  |  | MIL-DTL- 17111E§4.5.3.4 | 48
|  | FTM 3403 | rcf 600-700 | Annual |  | 2016-03-11 | SR | 2017-03-11 |  |  |  |  | 3403.2 | 49
|  | D91 | rcf 600-700 | Annual |  | 2016-03-11 | SR | 2017-03-11 |  |  |  |  | D91-02(2012) | 50
|  | D2273 | rcf 600-700 | Annual |  | 2016-03-11 | SR | 2017-03-11 |  |  |  |  | D2273-08(2012) | 51
|  | D893 | rcf 600-700 | Annual |  | 2016-03-11 | SR | 2017-03-11 |  | 2016-09-12 |  | 4000.20.110 | D893-14 | 52
| NACE | TM 0172 | Speed of chuck on grinder 1700-1800 rpm | annually |  | 2016-04-15 | SR | 2017-04-15 |  | 2014-04-17 | 2017-04-17 | 4000.20.083 4/14/14 |  | 53
|  |  | Diameter of specimen 9.53-12.7mm | quarterly |  | 2016-09-01 | JB | 2016-12-01 | [5 empty] | 54
|  |  | Temperature of bath 100 + 2 F | daily | 13-X-4 | with use | [7 empty] | 55
|  |  | Stirrer apparatus speed 1000 + 50 rpm | annually | 13-X-4 | 2016-04-15 | SR | 2017-04-15 | [5 empty] | 56
| Freeze Point | D2386 | Stirrer 1-1.5 cycle/s | annually | N/A | 2016-04-08 | SC/AC | 2017-04-08 |  | 2014-09-23 | 2015-09-23 | 4000.20.007 6/14/11 | D2386-06(2012) | 57
|  | D1177 | Stirrer 60-80 strokes per minute | annually | [9 empty] | 58
| Pour Point | D97 | verify bath temp 0 + 1.5C | Daily | [5 empty] | 2013-07-03 | Method updated | 4000.20.027 7/24/13 | D97-16 | 59
|  |  | verify bath temp -18 + 1.5C | Daily | [9 empty] | 60
|  |  | verify bath temp -33 + 1.5C | Daily | [9 empty] | 61
|  |  | verify bath temp -51 + 1.5C | Daily | [9 empty] | 62
|  |  | verify bath temp -69 + 1.5C | Daily | [9 empty] | 63
| Cloud Point | D2500 | verify bath temp 0 + 1.5C | Daily | [5 empty] | 2012-12-12 | Method updated | 4000.20.026 | D2500-16 | 64
|  |  | verify bath temp -18 + 1.5C | Daily | [9 empty] | 65
|  |  | verify bath temp -33 + 1.5C | Daily | [9 empty] | 66
|  |  | verify bath temp -51 + 1.5C | Daily | [9 empty] | 67
|  |  | verify bath temp -69 + 1.5C | Daily | [9 empty] | 68
| Koehler Cloud/Pour bath |  | Verify temperature of each controller annually | annually | R82590067 | 2016-08-02 | JRW | 2017-08-02 | [5 empty] | 69
| Durometer | D471 |  | Annually | 913624 (811-344) | 2016-11-11 | Micro Precision | 2017-11-11 |  |  |  |  | D471-16 | 70
| Durometer | D471 |  | Annually | 893792 (HH-332) | 2016-11-11 | Micro Precision | 2017-11-11 | [5 empty] | 71
| Bath | D3427 | Verify Temps at 25,50,75C | Annually | 82G-9981-X-1 | 2016-05-11 | MCA | 2017-05-11 | [5 empty] | 72
| Air Pressure | D3427 | Pressure reading | Annually | 280807 | 2016-05-11 | MCA | 2017-05-11 | [5 empty] | 73
| Heater Tube | D3427 | Verify Temps at 25,50,75C | Annually | 001X0X0 | 2016-05-13 | MCA | 2017-05-13 | [5 empty] | 74
| Oil Separation Pressure Gauge |  | verify accuracy with water 0.25psi-0.01psi | Annually |  | 2016-06-17 | MS | 2017-06-17 | [5 empty] | 75
| Penetrometer | D217 | Calibrate penetration dial | Annually | Labline | 2016-06-02 | MCA | 2017-06-02 | [5 empty] | 76
| Penetrometer | D1403 | Calibrate weight of shaft plus 1/4" cone | Annually | Precision Scientific 13AJ-1 | 2016-08-24 | MCA | 2017-08-24 | [5 empty] | 77
|  |  | Calibrate penetration distance in 1/10mm | Annually | Precision Scientific 13AJ-1 | 2016-08-24 | MCA | 2017-08-24 | [5 empty] | 78
| Water Seperability Bath | D1401 | Calibrate Stirring Paddle 1500+/-15rpm | Annually | Precision Scientific LM-395 | 2016-06-03 | MCA | 2017-06-03 | [5 empty] | 79
|  |  | Calibrate Bath @54C | Annually | Precision Scientific LM-395 | 2016-06-03 | MCA | 2017-06-03 | [5 empty] | 80
|  |  | Calibrate Bath @82C | Annually | Precision Scientific LM-395 | 2016-06-03 | MCA | 2017-06-03 | [5 empty] | 81
| Life Performance of Automotive Wheel Bearing Grease | D4290/D3527 | Motor RPM (1000) | Annually | Koehler K-18500 | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 82
| Life Performance of Automotive Wheel Bearing Grease | D4290/D3527 | Chamber Temp. | Annually | Koehler K-18500 | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 83
| Life Performance of Automotive Wheel Bearing Grease | D4290/D3527 | Spindle Temp. | Annually | Koehler K-18500 | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 84
| Roll Stability | D1831 | Roller RPM | 6-months | Precision Scientific 73437 | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 85
| Roll Stability | D1831 | Heater | 6-months | Precision Scientific 73437 | 2016-09-27 | MCA | 2017-09-27 | [5 empty] | 86
| Low Temp. Torque | D4693 | Time one revolution of spindle to 60sec to calibrate 1rpm | 2 years | Phoenix Chem Lab | 2016-10-11 | MCA | 2018-10-11 | [5 empty] | 87
| Low Temp. Torque | D4693 | Load cell output in meter and data logger | 6-months | Phoenix Chem Lab | 2016-10-11 | MCA | 2017-04-11 | [5 empty] | 88
| Data Logger for Low Temp Torque | D4693 |  | 6-Months | MC Measurement Computing USB-2416 | 2016-10-11 | MCA | 2017-04-11 | [5 empty] | 89
| Torque Meter for Low Temp Torque | D4693 | Use tench command torque valves in N.m with 100mm distance from centerline | 6-months | Automation Direct Prosense DPM3-AT-A-H | 2016-10-11 | MCA | 2017-04-11 | [5 empty] | 90
| Brookfield Viscosity |  | Calibrate each spindle to a known reference fluid within the range of the spindle | Annually | RVF 4124 | 2016-08-23 | MCA | 2017-08-23 | [5 empty] | 91
| Standard Deposition |  | temperature +/1C | Annually |  | 2016-11-14 | MS | 2017-11-14 | [5 empty] | 92
| Standard Deposition |  | Pressure +/- 2% | Annually |  | 2016-11-14 | MS | 2017-11-14 | [5 empty] | 93
| Standard Deposition |  | Air Flow +/- 10% | Annually |  | 2016-11-14 | MS | 2017-11-14 | [5 empty] | 94

## Octane Room
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | Comments | TODO ID
Octane Room | Balance |  | Calibrate | every 6 months | OHAUS CS5000 | Taken out of service 3/16/15 | SR | [6 empty] | Taken out of service 3/16/15 | 1
|  |  | Calibrate | every 6 months | OHAUS SP2000 | 2016-02-02 | Aldinger | 2016-08-02 | [6 empty] | 2
Octane Room | RON | D2699 | [7 empty] | 2013-08-13 | 2016-08-13 | 4000.20.019 12/20/12 | D2699-13b |  | 3
| MON | D2700 | [7 empty] | 2013-08-12 | 2016-08-13 | 4000.20.028 1/30/14 | D2700-14 |  | 4

## Viscometer Constants
Calibration, Verification, Standardization Schedule
Size | Serial # | Constant | Date Calibrated | Analyst | Annual Verification Date | Analyst | Calibration due date | Section J | TODO ID | NOT DATA
Semi-Micro Cannon-Manning | [9 empty] | 1
75 | C396 | @40=0.008283, | 2016-10-20 | CE |  |  | 2021-10-20 |  | 1
75 | C398 | @40=0.00863, @100=0.00861 | [6 empty] | 2
75 | C399 | @40=0.00743 | 2016-10-20 | CE |  |  | 2021-10-20 |  | 3
75 | E62 | @40C-0.0099559 | 2015-07-15 | SC |  |  | 2020-07-15 |  | 4
100 | A945 | @40=-.01421, @100=0.01418 | [6 empty] | 5
100 | A967 | @40C- 0.01449 | 2015-09-29 | SC |  |  | 2020-09-29 |  | 6
100 | C640 | @40C- 0.01590 | 2016-10-13 | CE |  |  | 2021-10-13 |  | 7
100 | C722 | @40C-0.01351 | 2016-10-13 | CE |  |  | 2021-10-13 |  | 8
100 | C723 | @40C=0.01447 | 2016-10-24 | CE |  |  | 2021-10-24 |  | 9
150 | A586 | 0.034465 | [6 empty] | 10
150 | A608 | @40C-0.034256 | 2015-09-29 | SC |  |  | 2020-09-29 |  | 11
150 | C666 | @40-0.03798, @100-0.03792 | 2016-11-21 | Cannon |  |  | 2021-11-21 |  | 12
150 | C667 | @40-0.02996, @100-0.02990 | 2016-11-21 | Cannon |  |  | 2021-11-21 |  | 13
200 | A379 | 0.09657 | [6 empty] | 14
200 | A499 | 0.09401 | 2016-02-29 | ME |  |  |  |  | 15
200 | A840 | 0.1004 | 2016-02-29 | ME |  |  |  |  | 16
200 | A935 | [7 empty] | 17
200 | B39 | 0.0906 | 2016-02-29 | ME |  |  |  |  | 18
200 | c374 | @40=0.0977, @100=0.0975 | [6 empty] | 19
200 | C375 | @40=0.1010- needs checked, @100=0.1008 | @ 100C 9/22/2015 | KAR |  |  |  |  | 20
200 | C376 | @40=0.09526, @100=0.09344 -needs checked | @40C 9/30/2016 @100c 9/22/2015 | @40C JS/MM KAR |  |  | 2017-09-30 |  | 21
300 | A670 | @40C-0.2455 | 2014-10-09 | KAR |  |  | 2019-10-09 |  | 22
300 | A673 | 0.2294 | [6 empty] | 23
300 | A679 | @40C-0.2458 | 2015-01-15 | AG |  |  | 2020-01-05 |  | 24
300 | A680 | @40=0.2392, @100=0.2385 | [6 empty] | 25
300 | A793 | @40C-0.24846 | 2015-07-14 | SC |  |  | 2020-07-14 |  | 26
400 | A274 | @40C-1.14015 | 2014-10-09 | KAR |  |  | 2019-10-09 |  | 27
400 | A275 | [7 empty] | 28
Ubbelohde Low-Temp | [9 empty] | 1
1 | 7072 | 0.008721 | 2015-07-17 | SC |  |  | 2016-07-17 |  | 29
1 | F37 | 0.01 | [6 empty] | 30
1 | I545 | 0.009603 | 2014-10-09 | KAR |  |  | 2015-10-09 |  | 31
1 | M346 | 0.009426 | 2016-02-15 | DA |  |  | 2017-02-15 |  | 32
1 | M204 | 0.009498 | 2016-02-15 | DA |  |  | 2017-02-15 |  | 33
1C | F899 | 0.02978 | [6 empty] | 34
1C | F977 | 0.03143 | 2015-08-31 | SC/SR |  |  | 2020-08-31 |  | 35
2 | 574 | 0.09688 | 2016-11-15 | SC |  |  | 2021-11-15 |  | 36
2 | G597 | 0.1006 | 2015-03-04 | AG |  |  | 2020-03-04 |  | 37
2B | M352 | 0.5626 | 2015-01-22 | AG |  |  | 2020-01-22 |  | 38
2C | M343 | 0.3206 | 2015-01-22 | AG |  |  | 2020-01-22 |  | 39
3B | 176 | 4.707 | 2016-05-31 | SC |  |  | 2021-05-31 |  | 40
3B | 177 | 4.8758 | 2016-05-31 | SC |  |  | 2021-05-31 |  | 41
3B | 621 | 4.4488 | 2015-01-22 | AG |  |  | 2020-01-22 |  | 42
3B | A262 | 4.8342 | 2014-10-16 | SA |  |  | 2019-10-16 |  | 43
3B | K557 | 4.5817 | 2016-11-15 | SC |  |  | 2021-11-15 |  | 44
3B | K570 | 4.667 | 2016-11-15 | SC |  |  | 2021-11-15 |  | 45
3B | K572 | 4.854 | 2015-02-26 | AG |  |  | 2020-02-26 |  | 46
3C | L95 | 2.862 | 2015-04-08 | AG |  |  | 2020-04-08 |  | 47
3C | N114 | 3.048 | 2016-08-30 | Cannon |  |  | 2021-08-30 |  | 48
4 | 1609 | 11.6984 | 2015-07-15 | SC |  |  | 2020-07-15 |  | 49
4 | 1630 | 9.5 | [6 empty] | 50
4 | L71 | 11.8415 | 2015-07-15 | SC |  |  | 2020-07-15 |  | 51
4 | L91 | 9.43 | 2012-08-23 | Cannon |  |  | 2017-08-23 |  | 52
4 | L98 | 10.9788 | 2015-07-14 | SC |  |  | 2019-07-14 |  | 53
4B | 145 | 54.58 | 2016-11-03 | SC |  |  | 2017-11-03 |  | 54
4B | A244 | 60.55 | [6 empty] | 55
4B | A566 | 64.34 | 2015-01-20 | AG |  |  | 2016-01-20 |  | 56
4B | C482 | 50.85 | 2016-11-04 | SC |  |  | 2017-11-04 |  | 57
4B | C484 | 50.56 | [6 empty] | 58
4B | C494 | 52.24 | 2016-11-03 | SC |  |  | 2017-11-03 |  | 59
4B | C506 | 51.9 | 2016-04-01 | JL |  |  | 2017-04-01 |  | 60
4B | C507 | 51.7 | 2016-11-03 | SC |  |  | 2017-11-03 |  | 61
4B | C512 | 52.11 | [6 empty] | 62
4B | C525 | 52.19 | 2016-11-04 | SC |  |  | 2017-11-04 |  | 63
4B | C526 | 51.65 | 2016-04-01 | JL |  |  | 2017-04-01 |  | 64
4B | C535 | 51.52 | 2016-04-01 | JL |  |  | 2017-04-01 |  | 65
4C | 713 | 30.15 | [6 empty] | 66
4C | A99 | 27.84 | [6 empty] | 67
4C | A110 | 31.59 | [6 empty] | 68
4C | A120 | 30.13 | [6 empty] | 69
4C | A133 | 28.74 | [6 empty] | 70
4C | A192 | 30.64 | 2015-01-20 | AG |  |  | 2020-01-20 |  | 71
4C | A447 | 28.34 | [6 empty] | 72
4C | C671 | 33.29 | 2015-07-28 | KAR |  |  | 2020-07-28 |  | 73
4C | C686 | 33.59 | 2015-01-20 | AG |  |  | 2020-01-20 |  | 74
4C | C745 | 30.97 | 2015-01-21 | AG |  |  | 2020-01-21 |  | 75
4C | C747 | 32.45 | 2015-01-20 | AG |  |  | 2020-01-20 |  | 76
5 | A667 | 129.277 | 2015-07-20 | KAR |  |  | 2016-07-20 |  | 77
5 | K531 | 103.9 | 2016-11-01 | SC |  |  | 2017-11-01 |  | 78
5 | K549 | 105.2 | 2016-11-01 | SC |  |  | 2017-11-01 |  | 79
5 | K597 | 103.9 | 2016-11-01 | SC |  |  | 2017-11-01 |  | 80
5 | K598 | 105.1 | 2016-11-01 | SC |  |  | 2015-08-20 |  | 81
Cannon-Fenske | [9 empty] | 1
25 | H332 | @40C-0.001724 | 2015-02-26 | AG |  |  | 2020-02-26 |  | 82
25 | A679 | @40C=0.002101 | 2016-11-08 | SB |  |  | 2021-11-08 |  | 83
50 | 1602 | @40C=0.004203 | 2016-10-24 | CE |  |  | 2021-10-24 |  | 84
50 | 13A | 0.004017@40C, 0.003999@ 100C | 2015-04-01 | Cannon |  |  | 2020-04-01 |  | 85
50 | 29A | 0.003902@40C, 0.003884 @ 100C | 2015-04-01 | Cannon |  |  | 2020-04-01 |  | 86
50 | 28A | 0.004080@ 40C, 0.004062 @ 100C | 2015-04-01 | Cannon |  |  | 2020-04-01 |  | 87
50 | 31A | 0.004176@ 40C, 0.004158@ 100C | 2015-04-01 | Cannon |  |  | 2020-04-01 |  | 88
50 | M624 | @40=0.004284, @100=0.004267 | [6 empty] | 89
50 | U679 | @40C=0.003930 | 2016-10-14 | CE |  |  | 2021-10-14 |  | 90
75 | 638 | @40- 0.007925 | 2016-01-15 | DA |  |  | 2021-01-15 |  | 91
75 | A512 | @40C-0.007456 | 2014-10-16 | SA |  |  | 2019-10-16 |  | 92
75 | A515 | @40C=0.007588 | 2016-10-20 | CE |  |  | 2021-10-20 |  | 93
75 | 8966 | @40=0.008309 | 2016-01-15 | DA |  |  | 2021-01-15 |  | 94
75 | 8967 | [7 empty] | 95
100 | 1560S | @40C- 0.015455 | 2014-10-07 | KAR |  |  | 2015-10-07 |  | 96
100 | 3208S | 0.01528 @ 40 | 2015-10-25 | KAR |  |  | 2016-10-25 |  | 97
100 | 4762S | 0.01601 @ 40, 0.01603@100 C | 2016-11-15 | JS/MM |  |  | 2017-11-15 |  | 98
100 | 4763S | , 0.01577 @ 40 C, 0.01570@100C | 2015-11-11 | DA |  |  | 2016-11-11 |  | 99
100 | 5620S | 0.01496-40C, 0.01516-100C | 2014-10-07 | KAR |  |  | 2015-10-07 |  | 100
100 | 9432S | [7 empty] | 101
100 | B115 | 0.01443 @ 40 | 2015-10-25 | KAR |  |  | 2016-10-25 |  | 102
100 | H1570 | 0.01529 @ 40; 0.01540 @ 100 | 2015-10-25 | KAR |  |  | 2016-10-25 |  | 103
100 | L581 | '0.01657@40C, 0.01666@100C | 2016-11-15 | JS/MM |  |  | 2017-11-15 |  | 104
100 | N971 | 0.01372@ 40C | 2015-11-11 | DA |  |  | 2016-11-11 |  | 105
100 | W656 | 0.01661-40C, 0.01652-100C | 2015-03-04 | AG |  |  | 2016-03-04 |  | 106
150 | 1145 | 0.03446- 40C, 0.03448- 100C | 2016-07-26 | ME |  |  | 2021-07-26 |  | 107
150 | 1149 | 0.0355 @ 40C | 2015-12-08 | DA |  |  | 2016-12-08 |  | 108
150 | 6871 | 0.03419 @ 40C, 0.03418@100C | 2016-07-26 | ME |  |  | 2021-07-26 |  | 109
150 | 3205S | 0.03578@ 100C | 2016-07-26 | ME |  |  | 2021-07-26 |  | 110
150 | 9326S | 0.03720 @ 40C, 0.03716 @ 100C | 2016-07-26 | ME |  |  | 2021-07-26 |  | 111
150 | J826 | 0.03909@40C, 0.03893 @ 100C | 2016-07-26 | ME | Broken- 10/9/16 |  | 2021-07-26 |  | 112
150 | W261 | @40C- 0.03560, @100C- 0.03563 | 10/4/16, 10/3/2016 | JS/MM, JS/MM |  |  | 10/4/2021@40C, 10/3/2021 @100C |  | 113
150 | X443F | @40C- 0.03409 | 2016-10-11 | CE |  |  | 2021-10-11 |  | 114
200 | 4764S | 0.1163@40C | 2016-11-16 | JS/MM |  |  | 2017-11-16 |  | 115
200 | 4766S | @40= 0.1195 | 2016-03-31 | JL |  |  | 2021-03-31 |  | 116
200 | 4767S | @40C- 0.1219 | 2016-09-30 | JS/MM |  |  | 2021-09-30 |  | 117
200 | B1449 | [7 empty] | 118
200 | D482 | [7 empty] | 119
200 | H70 | 0.1008@40C | 2016-11-16 | JS/MM |  |  | 2017-11-16 |  | 120
200 | H73 | 0.10305- 100C | 2016-07-26 | ME |  |  | 2021-07-26 |  | 121
200 | S861 | @40C-0.1016 | 2016-11-16 | JS/MM |  |  | 2017-11-16 |  | 122
200 | V915 | 0.09125@40C | 2016-11-16 | JS/MM |  |  | 2017-11-16 |  | 123
200 | V916 | 0.08929-40C, 0.08888-100C | [6 empty] | 124
350 | 598 | 0.4610 -40C | 2016-01-04 | NL |  |  | 2021-01-04 |  | 125
400 | 724 | @40=1.3270 | 2015-12-08 | DA |  |  | 2020-12-08 |  | 126
Cannon-Fenske Opaque | [9 empty] | 1
100 | A234 | @40=0.01585, @100=0.01595 | [5 empty] | @40=0.01151, @100=0.01160 | 127
150 | A157 | [7 empty] | 128
200 | A445 | @40=0.1018, @100=.1024 | [5 empty] | @40=0.0760, @100=0.0766 | 129
200 | B549 | [7 empty] | 130
300 | A463 | [7 empty] | 131
300 | B29 | [7 empty] | 132
350 | A193 | @40=0.5000, @100=0.5030 | [5 empty] | @40=0.3470, 2100=0.3498 | 133
400 | 569 | [7 empty] | 134
400 | F570 | 1.336 | 2015-02-09 | AG |  |  | 2020-02-09 | 0.947 | 135
450 | 337 | @40=2.444, @100=2.459 | [5 empty] | @40=1.707, @100=1.721 | 136
500 | 138 | [7 empty] | 137
500 | 644A | 7.452 | 2015-02-09 | AG |  |  | 2020-02-09 | 5.849 | 138
700 | C321 | [7 empty] | 139
700 | A806 | 129.4 | 2015-02-09 | AG |  |  | 2020-02-09 | 82.62 | 140
Zeitfuchs Cross-Arm | [9 empty] | 1
4 | C34 | [7 empty] | 141
Automated Viscometer | [8 empty] | 142
Right |  | 0.009672 | 2014-10-23 | JL |  |  |  |  | 143
4 | C35 | [7 empty] | 144
4 | C36 | [7 empty] | 145
4 | C38 | [7 empty] | 146
4 | C39 | [7 empty] | 147
4 | C40 | [7 empty] | 148
4 | C41 | [7 empty] | 149
4 | C42 | [7 empty] | 150
4 | C43 | [7 empty] | 151
4 | C44 | [7 empty] | 152
6 | B159Z | [7 empty] | 153
6 | B160Z | [7 empty] | 154
6 | B161Z | [7 empty] | 155
6 | B163Z | [7 empty] | 156
6 | B163Z | [7 empty] | 157
6 | B164Z | [7 empty] | 158
6 | B165Z | [7 empty] | 159
6 | B166Z | [7 empty] | 160
6 | B167Z | [7 empty] | 161
6 | B168Z | [7 empty] | 162

## Thermometers
Calibration, Verification, Standardization Schedule
Thermometer | Serial Number | Date Calibrated | Correction (amount that thermometer reading needs to be adjusted to match calibrator) | Next Due | Comments | TODO ID
ASTM 1F | 641431 | 2015-08-24 | 0 at 213F | 2016-02-24 | Not in use | 1
ASTM 4C Acid | 701102698 | 2016-02-05 | +1 at 100C, 0 at 40C, -1 at 0C | 2016-08-05 |  | 2
ASTM 5C | BB07503 | 2016-07-28 | None | 2017-01-28 |  | 3
ASTM 5C | BB07507 | 2015-08-17 | 0 at 2.3C, 0 at 44.3C | 2016-02-17 | Not in use | 4
ASTM 5C | BB07488 | 2015-08-17 | 0 at 2.2C, 0 at 44.3C | 2016-02-17 | Not in use | 5
ASTM 5C | BB07473 | 2015-08-17 | +0.4 at 2.4C, 0 at 44.3C | 2016-02-17 | Not in use | 6
ASTM 5C | BB06066 | 2015-08-17 | 0 at 4.3C, 0 at 44.3C | 2016-02-17 | Not in use | 7
ASTM 5C | BB007465 | 2016-07-28 | None | 2017-01-28 |  | 8
ASTM 5C | A07690 | 2015-08-17 | 0 at 4.6C, 0 at 44.3C | 2016-02-17 | Not in use | 9
ASTM 6C | A93-154 | 2016-09-01 | -0.5 at 0C, -1 at -20C, +4.5at -40C, 0 at -60C | 2017-03-01 |  | 10
ASTM 6C-86 | 778100 | 2016-02-01 | +3.5 at -20C, +3.0 at -40C, +3.5C at -80C | 2016-08-01 |  | 11
ASTM 6C-86 | 778105 | 2016-02-01 | 0 at 0C, 0C at -20C, +4C at -40C, +5C at -60C | 2016-08-01 |  | 12
ASTM 6C-86 | 778121 | 2016-02-01 | +1 at 0C, 0 at -20C, +4 at -40C, +5 at -60C | 2016-08-01 |  | 13
ASTM 8C | 777324 | 2015-03-16 | 0 | 2015-09-16 | Not in use | 14
ASTM 9F |  | 2014-09-21 | 1.401 | 2015-03-21 | Not in use | 15
ASTM 11C | 779942 | 2015-12-02 |  | 2016-06-02 |  | 16
ASTM 18F | 72902F |  |  |  |  | 17
ASTM 18F | 167677 | 2015-08-29 | 0 at 104.7F | 2016-02-19 | Not in use | 18
ASTM 18F | C622405 | 2015-08-19 | 0 at 105.3F | 2016-02-19 | Not in use | 19
ASTM 18F | 44407 | 2015-08-19 | -0.2 at 105.4F | 2016-02-19 | Not in use | 20
ASTM 18F | 423364 | 2015-08-19 | -0.2 at 105.4F | 2016-02-19 | Not in use | 21
ASTM 18F | 672444 | 2016-02-05 | 0 at 104F | 2016-08-05 |  | 22
ASTM 18F | 757-190 | 2016-02-01 | 0 at 104F | 2016-08-01 |  | 23
ASTM 18F | 42088 | 2015-08-19 | 0 at 105.0F | 2016-02-19 | Not in use | 24
ASTM 22F | K693991 | 2016-02-05 | +0.1 at212F | 2016-08-05 |  | 25
ASTM 22F | 76680 | 2016-02-05 | 0 at 212F | 2016-08-05 |  | 26
ASTM 34C | H67368 | 2015-08-04 | -0.3C at 100C | 2016-02-04 | Not in use | 27
ASTM 34C | T913270 | 2016-02-05 | -0.1 at 100C, 0 at 40C | 2016-08-05 |  | 28
ASTM 57C | 2760 | 2015-08-17 | -0.6 at 3.0C, 0 at 44.3C | 2016-02-17 | Not in use | 29
ASTM 62F | 19238 | 2015-08-17 | -0.2 at -3.8f | 2016-02-17 | Not in use | 30
ASTM 62F | 9147296 | 2015-08-17 | 0 at -4.1F | 2016-02-17 | Not in use | 31
ASTM 64C | 21075 | 2015-08-19 | 0 at 38.1C | 2016-02-19 | Not in use | 32
ASTM 64C | 9731810 | 2015-08-19 | 0 at 104.8F | 2016-02-19 | This thermometer is missing | 33
ASTM 64F | 110507 | 2015-08-19 | 0 at 105.0F | 2016-02-19 | Not in use | 34
ASTM 64F | 407581 | 2015-08-19 | 0 at 104.9F | 2016-02-19 | Not in use | 35
ASTM 64F | 12-3015 |  |  |  |  | 36
ASTM 67F | 24-2206 | 2015-08-24 | +1 at 212F | 2016-02-24 | Not in use | 37
ASTM 83F | 3C2871 | 2015-08-19 | -1 at 110F | 2016-02-19 | Not in use | 38
ASTM 83F | 3C2866 | 2016-02-05 | 0 at 104F | 2016-08-05 |  | 39
ASTM 83F | 3C2868 | 2015-08-19 | 0 at 109F | 2016-02-19 | Not in use | 40
ASTM 83F | 3C2867 | 2015-08-19 | 0 at 108F | 2016-02-19 | Not in use | 41
ASTM 83F | 2K4975 | 2015-08-19 | 0 at 108F | 2016-02-19 | Not in use | 42
ASTM 86F | 1604 | 2016-02-05 | 0 at 212F | 2016-08-05 |  | 43
ASTM 86F | 514351 | 2015-08-24 | +3 at 210F | 2016-02-24 | Not in use | 44
ASTM 88C | 2143 | 2015-07-21 | -0.4 at 100C | 2016-01-21 | Not in use | 45
ASTM 89C | 62912 | 2015-08-17 | +0.8 at 1.7C, +0.3 at -20.4 | 2016-02-17 | Not in use | 46
ASTM 88C | 2143 | 2015-07-21 | -0.4 at 100C | 2016-01-21 | Broken 7/15/15 | 47
[5 empty] | Taken out of service | 48
[5 empty] | Taken out of service | 49
ASTM 89C | 73177 | 2015-08-17 | +0..5 at 2.0, 0 at -20.0C | 2016-02-17 | Not in use | 50
ASTM 89C | 73201 | 2015-08-17 | +1.6 at 0.4C, +1 at -21.0C | 2016-02-17 | Not in use | 51
ASTM 91C | 2007747 | 2015-08-19 | 0 at 38.2C | 2016-02-19 | Not in use | 52
ASTM 91C | 2007740 | 2015-08-19 | 0 at 38.3C | 2016-02-19 | Not in use | 53
[5 empty] | Could not locate this thermometer on 8/4/15 | 54
ASTM 114C | 586978 | 2016-02-01 | -1.5 at 0C, -1 at -20C, +0.5 at -40C, 0 at -60C | 2016-08-01 |  | 55
ASTM 127C | 1561 | 2015-08-04 | -0.07C | 2016-02-04 | Not in use | 56
ASTM Saybolt | 2153368 | 2015-08-19 | -0.2 at 105.4F | 2016-02-19 | Not in use | 57
ASTM Viscos | 10056 | 2015-08-19 | -0.3 at 105.5F | 2016-02-19 | Not in use | 58
Cloudpoint | 9174ZZZ | 2016-07-28 | None | 2017-01-28 |  | 59
Eintauchtiefe-Immersion | N16B | 2016-11-08 | +3.0C at 100C, +3.0 at 150C | 2017-05-08 |  | 60
Freeze Point | 6339 | 2016-02-01 | -1 at 0, -2 at -20C, 0 at -40C, -4 at -60C | 2016-08-01 |  | 61
Freeze Point | 6251 | 2016-02-01 | -2 at 0, -2 at -20C, 0 at -40C, +1 at -60C | 2016-08-01 |  | 62
Long Stem | 150151080 | 2015-03-04 |  | 2017-03-04 |  | 63
Long Stem | 150151111 | 2015-03-04 |  | 2017-03-04 |  | 64
Long Stem | 150151073 | 2015-03-04 |  | 2017-03-04 |  | 65
Long Stem | 150151093 | 2015-03-04 |  | 2017-03-04 |  | 66
Long Stem | 150151107 | 2015-03-04 |  | 2017-03-04 |  | 67
Long Stem | 150151112 | 2015-03-04 |  | 2017-03-04 |  | 68
Long Stem | 150151106 | 2015-03-04 |  | 2017-03-04 |  | 69
Long Stem | 150151100 | 2015-03-04 |  | 2017-03-04 |  | 70
Long Stem | 150151146 | 2015-03-04 |  | 2017-03-04 |  | 71
Long Stem | 150151145 | 2015-03-04 |  | 2017-03-04 |  | 72
Long Stem for RVP |  | 2014-09-21 | 0 at 0C, 0.34 at 100C | 2015-03-21 | Not in use | 73
Mastech MS6514 | MBFHC 14377 | 2015-12-31 |  | 2016-01-04 | Not in use | 74
Mastech MS6514 | MBFHC 14365 | 2015-12-31 |  | 2016-01-04 | Not in use | 75
Mastech MS6514 | MBFHC 14317 | 2015-12-31 |  | 2016-01-04 | Not in use | 76
NACE | 2409794 | 2015-08-04 | 0 | 2016-02-04 | Not in use | 77
SAMA FT40 |  | 2015-08-04 | none at 100C | 2016-02-04 | Not in use | 78
Thermocouple | 39330 | 2015-07-22 | 0 | 2016-01-22 |  | 79
Mastech MS6514 Thermocouple | 39332 | 2016-09-27 | 0 | 2017-03-22 |  | 80
Thermocouple | 39336 | 2015-07-22 | 0 | 2016-01-22 |  | 81
-70 to +60F | Curtain USA 3D 8874 | 2015-08-17 | -1.5 at 39.3F | 2016-02-17 | Not in use | 82
-70 to +75F | Curtain USA | 2015-08-17 | -2 at 42F | 2016-02-17 | Not in use | 83
-12 to 150C | Made in China White/rea | 2015-08-24 | -2 at 103C | 2016-02-24 | Not in use | 84
-10 to +1200 | Princo | 2015-08-24 | -1 at 214F | 2016-02-24 | Not in use | 85
0 - 125F | Brooklyn P-M Thermo. Co. | 2015-08-17 | 0 at 36F, 0 at 112F | 2016-02-17 | Not in use | 86
5 to 150F | GM MFG CO Tropical | 2015-08-17 | +1 at 39F, 0 at 112F | 2016-02-17 | Not in use | 87
30 to 300F GB | 46656 | 2015-08-24 | +1 at 212F | 2016-02-24 | Not in use | 88
40 to 200C | Napco | 2015-08-24 | 0 at 100.5C | 2016-02-24 | Not in use | 89
79.5 to 100.5 C Brooklyn | 31059 |  |  |  |  | 90
100C bath | POA061413B1 | 2015-02-10 | 0.4 | 2015-08-10 | Not in use | 91
250 to 350 F Taylor | 7A895964 |  |  |  |  | 92
ERDCO Seta Flash therm | 539823 | 2016-02-24 | none at 32F, None at 212F | 2016-08-24 |  | 93
ASTM 11C COC Flash | 12044 | 2016-03-03 | none | 2016-09-03 |  | 94
ASTM 11C COC Flash | 12076 | 2016-03-03 | none | 2016-09-03 |  | 95
ASTM 11C COC Flash | 12039 | 2016-03-03 | none | 2016-09-03 |  | 96
ASTM 7C | 1M5122 | 2016-08-24 | none at 0C, +1C at 100C | 2017-02-24 |  | 97
ASTM 11-C | 12044 | 2016-08-24 | none at 0C, -1 at 100C | 2017-02-24 |  | 98
ASTM 3C | 11911 | 2016-08-24 | none at 0C, -1 at 100C | 2017-02-24 |  | 99
ASTM 8C | 90960 | 2016-08-24 | none at 0C, none at 100C | 2017-02-24 |  | 100
Digital on Drying oven | T150331-I | 2016-09-14 |  | 2017-09-14 |  | 101
Digital on Drying oven | T150331-C | 2016-09-14 |  | 2017-09-14 |  | 102
Data-Logger Thermocouple | EL-USB-TC-LCD | 2016-09-27 |  | 2017-03-27 |  | 103

## Barometers
Calibration, Verification, Standardization Schedule
Location | Equipment | ASTM | Requirement | Frequency | Serial Number | Last Calibration Date | Associate | Next Calibration Date | Associate | TMA Date | Next TMA Date | SOP Number and Date | Current ASTM Version | TODO ID
| Barometer |  | NIST traceable | biannually | 151922933 | 2015-12-04 | Control Comp | 2017-12-04 | Control Comp |  |  |  |  | 1
[5 empty] | 150443486 | 2015-05-26 |  | 2017-05-26 | [5 empty] | 2
[5 empty] | 150443489 | 2015-05-26 |  | 2017-05-26 | [5 empty] | 3
[5 empty] | 150443479 | 2015-05-26 |  | 2017-05-26 | [5 empty] | 4

## Micropipettors
Calibration, Verification, Standardization Schedule
Pipettor | Serial Number | Date Calibrated | Next Due | Comments | TODO ID
Gilson | R12442H | 2015-07-13 | 2015-08-13 | Taken out of service 8/6/15 | 1
LuminUltra 100uL |  | 2016-11-14 | 2016-12-14 |  | 2
LuminUltra | 66742 | 2016-11-14 | 2016-12-14 |  | 3
Emcee 50uL |  | 2016-11-14 | 2016-12-14 |  | 4
25uL | H10563A | 2016-06-16 | 2016-07-16 | Could not locate 7/15/16 | 5
25uL | E11350L | 2016-11-14 | 2016-12-14 |  | 6
1000ul | JK05218 | 2016-10-18 | 2016-11-18 | Could not locate 11/14/16 | 7
1000ul | AB15024 | 2016-11-14 | 2016-12-14 |  | 8
50uL | T128416 | 2016-11-14 | 2016-12-14 |  | 9
50uL | G10829C | 2016-11-14 | 2016-12-14 |  | 10
250uL | S131356 | 2016-11-14 | 2016-12-14 |  | 11
250uL | Y12240D | 2016-11-14 | 2016-12-14 |  | 12
250uL | T12854G | 2016-11-14 | 2016-12-14 |  | 13
250uL | KB05101 | 2016-11-14 | 2016-12-14 |  | 14
10uL | W12571A | 2016-11-14 | 2016-12-14 |  | 15
Eppendorf 5000uL | K19246D | 2016-11-14 | 2016-12-14 |  | 16
Transferpettor 500uL | 03F96954 |  |  | Needs tips 6/16/16 | 17
Eppendorf 1000uL | H28124D | 2016-11-14 | 2016-12-14 |  | 18

## Timers
Calibration, Verification, Standardization Schedule
Timer | Serial Number | Date Calibrated | Next Due | Comments | TODO ID
Yellow Taylor Kitchen Timer | 150313-A | 2016-09-13 | 2017-09-13 |  | 1
Yellow Taylor Kitchen Timer | 150313-B | 2016-09-13 | 2017-09-13 |  | 2
Yellow Taylor Kitchen Timer | 150313-C | 2016-09-13 | 2017-09-13 |  | 3
Yellow Taylor Kitchen Timer | 150324-B | 2016-09-13 | 2017-09-13 |  | 4
Yellow Taylor Kitchen Timer | 150118-A | 2016-09-13 | 2017-09-13 |  | 5
Yellow Taylor Kitchen Timer | 150118-B | 2016-09-13 | 2017-09-13 |  | 6
Grey Kitchen Timer | 160913-A | 2016-09-13 | 2017-09-13 |  | 7
Grey Kitchen Timer | 160913-B | 2016-09-13 | 2017-09-13 |  | 8
Grey Kitchen Timer | 160913-C | 2016-09-13 | 2017-09-13 |  | 9
Grey Kitchen Timer | 160913-D | 2016-09-13 | 2017-09-13 |  | 10
Fisher Scientific | 140673378 |  | 2016-10-13 | Not found in Lab 9/13/16 | 11
Fisher Scientific | 140673385 |  | 2016-10-14 |  | 12
Fisher Scientific | 140673392 |  | 2016-10-13 | Not found in Lab 9/13/16 | 13
Fisher Scientific | 140673384 |  | 2016-10-13 | Not found in Lab 9/13/16 | 14
Fisher Scientific | 140673368 |  | 2016-10-13 | Not found in Lab 9/13/16 | 15
Fisher Scientific | 140673401 |  | 2016-10-13 | Not found in Lab 9/13/16 | 16
Fisher Scientific | 140673398 |  | 2016-10-13 | Not found in Lab 9/13/16 | 17
Fisher Scientific | 150084413 |  | 2017-02-05 | Not found in Lab 9/13/16 | 18
Fisher Scientific | 150084411 |  | 2017-02-05 | Not found in Lab 9/13/16 | 19
Fisher Scientific | 150084401 |  | 2017-02-05 | Not found in Lab 9/13/16 | 20
Fisher Scientific | 150084404 |  | 2017-02-05 | Not found in Lab 9/13/16 | 21
Thomas Scientific | 150501283 |  | 2017-07-01 | Not found in Lab 9/13/16 | 22
Thomas Scientific | 150501284 |  | 2017-07-01 | Not found in Lab 9/13/16 | 23
Thomas Scientific | 150501277 |  | 2017-07-01 |  | 24
Thomas Scientific | 150501286 |  | 2017-07-01 | Not found in Lab 9/13/16 | 25
Thomas Scientific | 140755808 |  | 2016-11-18 | Not found in Lab 9/13/16 | 26
Thomas Scientific | 150501279 |  | 2017-07-01 |  | 27
Thomas Scientific | 150501285 |  | 2017-07-01 |  | 28
Thomas Scientific | 150501280 |  | 2017-07-01 |  | 29
Thomas Scientific | 150501281 |  | 2017-07-01 |  | 30
Thomas Scientific | 151859815 |  | 2017-11-16 |  | 31
Thomas Scientific | 151859818 |  | 2017-11-16 |  | 32
Thomas Scientific | 160298478 |  | 2018-03-22 |  | 33
Thomas Scientific | 160298481 |  | 2018-03-22 |  | 34
Thomas Scientific | 160440291 |  | 2018-05-12 |  | 35
Thomas Scientific | 160440292 |  | 2018-05-12 |  | 36
Thomas Scientific | 160440293 |  | 2018-05-12 |  | 37
Thomas Scientific | 160440294 |  | 2018-05-12 | Not found in Lab 9/13/16 | 38
PRECISION Sientific Petroleum Instruments | 1447320 | 2015-06-24 | 2017-06-24 |  | 39
Thomas Scientific | 160612861 |  | 2018-07-21 |  | 40
Thomas Scientific | 160612863 |  | 2018-07-28 |  | 41
Thomas Scientific | 160612859 |  | 2018-07-21 |  | 42
Thomas Scientific | 160440295 |  | 2018-05-12 |  | 43
Thomas Scientific | 160612860 |  | 2018-07-21 |  | 44

## Balances
Calibration, Verification, Standardization Schedule
Balance Model # | Serial Number | Location | Date Calibrated | Calibrated by: | Next Due | TODO ID
HRB10001 | 1007 | Tribology Room | 2016-08-23 | Aldinger | 2017-02-23 | 1
AT201 | 1115323171 | Front Lab | 2016-08-23 | Aldinger | 2017-02-23 | 2
AP210 | 1118513346 | Front Lab | 2016-08-23 | Aldinger | 2017-02-23 | 3
HRB10001 | HR1409504 | Back Lab | 2016-08-23 | Aldinger | 2017-02-23 | 4
LF225DL | 138975036 | Back Lab | 2016-08-23 | Aldinger | 2017-02-23 | 5
AT20 | 1113250378 | Server Room | 2016-08-23 | Aldinger | 2017-02-23 | 6
SP2001 | B207693885 | Engine Room | 2016-08-23 | Aldinger | 2017-02-23 | 7